## Sheet1
Anesthetic Agent Analyzer
1 | DESCRIPTION
1.1 | Anaesthetic gas Analyzer is a tool designed for testing anaesthetic vaporizer output concentration.
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be a digital 5-agent gas analysis technology,must be as versatile as possible, portable, reliable, and compact.
2.2 | The unit offered must be capable of testing Anesthetic vaporizer output concentration in hospitals and outpatient surgical centers.
2.3 | The unit offered must automatically identifies and quantifies five anesthetic agents
2.4 | The unit offered must be capable to detect and quantifies mixed agents.
2.5 | The gas monitoring method of the unit offred should be non-dispersive infrared (NDIR) technology that identifies and measures five anesthetic agent gases: Halothane, Enflurane, Isoflurane,Desflurane, and Sevoflurane.
2.6 | The unit offered must also measures O2, CO2, and N2O
2.7 | Gas flow sampled at 0, 50, 100, 150 and 200 mL/min
2.8 | Fast warm-up time to ensure full accuracy within minutes
2.9 | The unit must have built in Auto-calibration
2.10 | Bright, real-time display of waveforms and numerical values provided
2.11 | The unit offered must be fitted with internal rechargeable batteries
2.12 | The Analyzer offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyzer may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified. The unit must come in a carry case.
2.13 | Must be provided with a manual guide preferably in the English language.