## Sheet1
Electrical Safety Analyzer
1 | DESCRIPTION
1.1 | An analyzer that tests the electrical output of medical devices.
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must provide one touch testing, which will allow the user to perform rapid safety tests on various medical devices, and must be very user friendly. The unit must provide electrical safety testing via automated (Onboard and/or PC Software driven) and manual modes for multiple safety standards tests that must meet IEC60601-1 and help the user to easily perform a complete electrical safety test in accordance with IEC60601-1 (Including 2nd and 3rd edition), IEC62353, VDE751 and ANSI/AAMI ES1 standards for medical equipment
2.2 | Bidder to state the method of automated testing (Onboard and/or PC software driven).
2.3 | The unit offered must provide user-programmable test sequences (Onboard and/or PC software driven)
2.4 | The unit must incorporate a built in ECG simulator, for ECG waveform simulation, with dual-lead measurement capability that combine the functionality of a simulator, multimeter and safety analyzer in a single test tool
2.5 | The Analyzer must have a reasonable size display for display of test values in the English Language and this display must be clearly visible under all lighting conditions.
2.6 | The unit must be able to automatically flag and warn of any parameter that is outside the allowable limit, based on the safety standard employed (Onboard and/or PC software driven)
2.7 | Upon completion of testing, the unit must be capable of storing test results, either onboard or via PC driven software, for data base storage and printing of test results for record keeping
2.8 | The Analyzer must be able to operate on 220 Volt + 10% 50hz single phase a.c. power supply. The Analyzer must be surge protected.
2.9 | The unit must be supplied with a suitable adapter of S.A.B.S. standard (15 amp 220 Volt a.c. power receptacle), which will accept 15 amp plug tops of the different equipment to be tested
2.1 | The unit must be supplied with a soft carry case and the unit must be fully portable
2.11 | Specification of the unit being tendered for must be as follows:
2.11.1 | Volts:
2.11.1.1 | Range: minimum range of 0 Volts to 300 Volts ac rms
2.11.2 | Insulation Resistance:
2.11.2.1 | Range: minimum range of 0,5 to 100 Mega Ohms
2.11.2.2 | Accuracy: the accuracy must not exceed a maximum of + 8% of reading
2.11.3 | Protective Earth Continuity and Test Current:
2.11.3.1 | P.E Range: typically 0.0 to 0,2 Ohms and testing current of 25A ac (IEC60601-1) and/or 200mA ac (IEC62353)
2.11.3.2 | Accuracy: accuracy must not exceed a maximum of + 5% of reading
2.11.4 | Equipment Current Consumption:
2.11.4.1 | Range: minimum range of 0 to 15 amp
2.11.4.2 | Accuracy: must not exceed a maximum of + 5% of reading
2.11.5 | Leakage Current / IEC60601-1; AAMI:
2.11.5.1 | Range: minimum range of 0 to 8,000uA
2.11.6 | Waveform Simulation
2.11.6.1 | Normal sinus rhythm:
2.11.6.1a | Minimum range selection between 30 and 240bpm
2.11.6.2 | Performance Waves:
2.11.6.2a | Pulse of typically 30 and 60bpm
2.11.6.2b | Sine wave of typically 10 to 100 Hz
2.11.6.2c | Square wave of typically (50% duty cycle) 0,125 and 2Hz
2.12 | Mains on Applied Part (MAP) must generate a signal of at least 100% (IEC60601-1 and IEC62353) of incoming mains power before applying it to the selected applied part. The signal must be applied with both normal and reversed polarity
2.13 | All internal electrical components must be protected against liquid leaks and liquid splashes
2.14 | The unit must have a communication port / connectivity (Wired or Wireless), which will enable connection to a PC for automation test software and firmware upgrades
2.14.1 | Bidder to provide the name/make of Automation Test Software as stated above at 2.14
2.15 | The automation test software, used with the unit, must be able to integrate with various Computerised Maintenance Management System (CMMS) software’s, which is an advantage for future implementation.
2.15.1 | Bidder to provide the name and make of various CMMS programs currently compatible with the unit
2.16 | The Analyser offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyzer may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified.
2.17 | Must be provided with a manual guide preferably in the English language.