## Sheet1
Incubator Analyzer
1 | DESCRIPTION
1.1 | Incubator analyzer is the analyzer that shows the performance of the incubator
2 | TECHNICAL SPECIFICATION
2.1 | The Incubator analyzer must be compatible and be suitable for testing and verifying all makes and models of incubators, both closed, forced-convection incubators, and open infant warmers, including transportable, air-controlled, and baby-controlled units, and must be supplied with a carry case
2.2 | The Analyzer offered must be as versatile as possible, portable and lightweight, and must fit inside the incubator and operate as a stand-alone device or with a personal computer for automated testing
2.3 | The Analyzer offered must be user friendly, with a menu key interface, allowing quick access to all tests, and in stand-alone mode, must display measured parameters repeatedly in a cycle fashion
2.4 | The Analyzer offered must be supplied with Windows compatible automation software for easy data collection, analysis, and documentation. The software must be supplied as part of the final offer at no extra cost
2.5 | The Analyzer offered must be designed around AAMI and IEC standards that specify incubator sound levels and thermal characteristics, and must accurately and simultaneously measure humidity, airflow, sound, and 4 independent temperatures at once, in a single, integrated unit
2.6 | The Analyzer offered must be able to perform up to 24-hours of continuous testing and monitoring with battery power only
2.7 | The Analyzer must be able to measure Temperature as follows:
2.7.1 | Measurement Range atleat 5 deg C to 50 deg C
2.7.2 | Accuracy: +/- 0.5 deg C
2.8 | The Analyzer must be able to measure Sound Level as follows:
2.8.1 | Measurement Range at least 30 dbA to 80 dbA
2.8.2 | Accuracy: +/- 5 dbA
2.9 | The Analyzer must be able to measure Relative Humidity as follows:
2.9.1 | Measurement Range: 0% to 100% RH, non-condensing
2.9.2 | Accuracy: +/- 6% RH
2.1 | The Analyzer must be able to measure Airflow as follows:
2.10.1 | Measurement Range at least 0.2m/s to 2.0m/s
2.10.2 | Accuracy: +/- 0.1 m/s
2.11 | The Analyzer offered must have an internal memory for storing measurements. Please state storage capacity. Must be able to store atleast 50 Test.
2.12 | The Analyzer must be able to operate on 220 Volt + 10% 50Hz single phase a.c. power supply. The Analyzer must be surge protected.
2.13 | Bidder to specify what the approximate present calibration cost on the offered Analyser will be, and the frequency of the calibration required for the tendered Analyzer
2.14 | The Analyzer offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyser may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified. The unit must come in a carry case.
2.15 | Must be provided with a manual guide preferably in the English language.