## Sheet1
Multi-Parameter Patient Simulator
1 | DESCRIPTION
1.1 | The multi-parameter simulator being tendered for must be of robust construction; it must be very user-friendly and simple to operate; it must comprise the latest technology; have a rechargeable battery pack.
2 | TECHNICAL SPECIFICATION
2.1 | The multi-parameter simulator must measure the following parameters:
2.1.1 | ECG
2.1.2 | Arrhythmias
2.1.3 | Pacemaker Waveforms
2.1.4 | Invasive Blood Pressures (IBP)
2.1.5 | Respiration
2.1.6 | Temperature
2.1.7 | Non-invasive Blood Pressure (NIBP)
2.1.8 | SPO2 (Oxygen Saturation/Pulse Rate)
2.2 | The ECG simulation on the unit being tendered for must be available as follows:
2.2.1 | A twelve (12) lead ECG simulation must be provided via 10 universal ECG jacks that are colour coded to IEC standards
2.2.2 | The heart rate must be user selectable in a minimum range of 20 to 300 BPM.
2.2.3 | The ECG amplitudes must be user selectable in the range 0,5mV to 5.0mV.
2.2.4 | ECG waveform selections must be provided from Adult to Paediatrics
2.2.5 | ECG performance must provide:
2.2.6 | Square wave of 2Hz and 0,125Hz.
2.2.7 | Pulse of at least 30 and 60 BPM with 60ms pulse width.
2.2.8 | Sine waves in a user selectable range of at least 0,5Hz to 60Hz.
2.2.9 | Triangle wave of at least 2Hz.
2.2.10 | ST segment elevation must be available in the range – 0,8mV to + 0,8mV in 0,1mV steps
2.2.11 | The ECG posts must be integrated into the simulator and allow for easy and secure 10-lead ECG Snap and Lead connections
2.3 | ARRHYTHMIA simulation selections must provide the following:
2.3.1 | Sinus Arrhythmia.
2.3.2 | Atrial (PAC).
2.3.3 | Atrial Tachycardia.
2.3.4 | Atrial Flutter.
2.3.5 | Premature Nodal Contraction (PNC).
2.3.6 | Supraventricular Tachycardia.
2.3.7 | PVC1 Left Ventricular Focus.
2.3.8 | PVC1 Early, LV Focus.
2.3.9 | PVC1 R on T, LV Focus.
2.3.10 | Missed Beat.
2.3.11 | PVC2 Right Ventricular Focus.
2.3.12 | PVC2 Early, RV Focus.
2.3.13 | PVC2 R on T, RV Focus.
2.3.14 | Multifocal PVC’s.
2.3.15 | Atrial Fibrillation Coarse / Fine.
2.3.16 | PVC’s (6, 12 or 24 per minute)
2.3.17 | Frequent Multifocal PVC’s.
2.3.18 | Bigeminy.
2.3.19 | Trigeminy.
2.3.20 | Multiple PVC’s (One time run of 2, 5 or 11 PVC’s)
2.3.21 | Ventricular Tachycardia.
2.3.22 | Ventricular Fibrillation Coarse / Fine.
2.3.23 | Asystole.
2.3.24 | Conduction Defect
2.4 | The pacemaker simulations must provide the following MINIMUM functions:
2.4.1 | Asynchronous Pacemaker
2.4.2 | Pacer Non Function.
2.4.3 | Pacer Non Capture.
2.4.4 | Demand Occasional Sinus.
2.4.5 | Demand Frequent Sinus.
2.4.6 | AV Sequential.
2.5 | A minimum of two Invasive Blood Pressure channels must be provided, and the unit being tendered for must be available as follows:
2.5.1 | The channels must provide a transducer sensitivity of 5 or 40 uV / V / mmHg
2.5.2 | Static Pressure levels for both IBP channels must be available in a stepped range of at least –5mmHg to at least +250mmHg in 1 mmHg steps
2.5.3 | Systolic and Diastolic pressure levels must be independently adjustable, in 1 mmHg steps, for any selected Dynamic Waveform
2.5.4 | Dynamic waveform selections must be provided as follows:
2.5.4a | Arterial: 120/80.
2.5.4b | Radial Artery: 120/80.
2.5.4c | Left Ventricle: 120/00.
2.5.4d | Right Ventricle: 25/00.
2.5.4e | Central Venous (CVP): 15/10.
2.5.4f | Pulmonary Artery: 25/10.
2.5.4g | Pulmonary Artery Wedge: 10/2
2.5.5 | The bidder must provide the catalogue number, the price inclusive of V.A.T. and the correct descriptions of all the available IBP interface cables for the different makes of monitors on a separate schedule
2.6 | Respiration simulation must provide the following:
2.6.1 | Baseline Impedance in a stepped range of at least 500 ohm to at least 1500 ohm for leads I, II and III must be provided
2.6.2 | The impedance variations must be typically in a range of: 1 to 5 Ohm
2.6.3 | The respiration rate must be available in a stepped range of at least 10 BrPM to at least 140 BrPM and 0 BrPM for apnea
2.6.4 | Apnea selections must be typically as follows:
2.6.4.1 | 12 seconds; 22 seconds; 32 seconds and continuous
2.7 | Temperature simulation must be available as follows:
2.7.1 | The temperature simulation must provide compatibility with both YSI Series 400 and 700
2.7.2 | The temperature simulation must be available in a stepped range of at least 30oC to at least 40oC
2.7.3 | Correct descriptions of all the available TEMP interface cables for the different makes of monitors on a separate schedule
2.8 | Non-Invasive Blood Pressure (NIBP) simulation must be available as follows:
2.8.1 | The NIBP Simulator must be suitable for testing and verifying all makes of Adult, Paediatric and Neonatal NIBP Monitors employing the oscillometric methods of measurement
2.8.2 | Dynamic Non-Invasive Blood Pressure:
2.8.2a | The unit must be able to simulate normal, hypertensive and hypotensive dynamic, non-invasive blood pressures and the simulations must represent typical Adult and Neonatal patients
2.8.2b | The unit must include an internal pump for Static calibration, Leak Tests and Overpressure tests
2.8.2c | The unit must deliver independent SYSTOLIC / DIASTOLIC pressure simulations and it must be possible to change these in steps of 1mmHg
2.8.3 | The unit must be able to provide Preprogrammed Target Value selections for SYSTOLIC / DIASTOLIC (MAP) Simulations typically as follows in mmHg:
2.8.3a | 60/30 (40)
2.8.3b | 80/50 (60)
2.8.3c | 100/65 (77)
2.8.3d | 120/80 (93)
2.8.3e | 150/100 (117)
2.8.3f | 200/150 (167)
2.8.3g | 255/195 (215)
2.8.4 | Bidder must state the selectable Target Values for NIBP simulations, if it differs from values above
2.8.5 | The NIBP Display Graphics must include the following:
2.8.5.1 | Dynamic Real-Time NIBP Cuff Pressure waveform
2.8.5.2 | The Displayed NIBP Real-Time Parameters must be summarised and include the following:
2.8.5.2a | Peak Cuff Pressure in mmHg
2.8.5.2b | Inflate / Deflate time in seconds
2.8.5.2c | Inflate / Deflate Rate in mmHg / second
2.8.5.2d | Total measurement Time in seconds
2.8.5.3 | The unit must have the ability to carry out SYSTEM LEAK TESTING, which must incorporate the following features:
2.8.5.4 | Adjustable Target Pressure setting, with a maximum pressure setting of 400mmHg
2.8.5.5 | System Leak Test result, to display the Leak Rate in mmHg / minute
2.8.5.6 | The unit must have the ability to carry out PRESSURE RELIEF TESTING, which must incorporate the following features:
2.8.5.7 | The unit must have the ability to test the operation of the monitor’s relief valve for a MAXIMUM PRESSURE of 400mmHg
2.8.5.8 | The cuff mandrels for attachment of the cuffs, for both ADULT and NEONATE, must be supplied as standard with the SIMULATOR on offer
2.8.5.9 | Compatible cuff adapters and tubing, for connecting the NIBP device to the simulator, must be available and supplied at no extra charge to the final tender price
2.9 | Non-Invasive Pulse Oximeter simulation must be available as follows:
2.9.1 | The unit offered provides Oximeter SpO2 optical emitter and detector capability (artificial finger), which generates an optical signal to verify that the electronics within the pulse oximeter probe are functional and presents pulse oximeter equipment with a signal, so that the operator can observe the resulting displayed value of SpO2, and compare it to the expected values
2.9.2 | The unit must be capable of simulating / testing the following manufacturers of Pulse Oximeters: BCI, Datex, Nellcor, Nihon Kohden, Nonin, Ohmeda, Masimo technology, Hewlett Packard, Philips, Mindray and GE
2.9.3 | The unit must be upgradeable in the future, should new makes of non-invasive Pulse Oximeters be introduced, in order to make testing of them possible
2.9.4 | The unit must have Preset R-curves for the makes and models of Pulse Oximeters mentioned above
2.9.5 | Bidder to state if Auto sequences are available for performance testing of SPO2 parameters
2.9.6 | Oxygen Saturation (SpO2) range of 30 – 100% with a resolution of 1%.
2.9.7 | Pulse / Heart rate in a minimum range of 30 – 250 BPM with a resolution of 1 BPM and accuracy of + 1% of setting. Bidder must state the accuracy
2.1 | The Simulator must have a reasonable size display for display of test values in the English Language and this display must be clearly visible under all lighting conditions.
2.11 | The unit must be portable, lightweight and have a protective carrying case.
2.12 | Bidder must state the type of re-chargeable battery employed and the average operating time
2.14 | The Simulator offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Simulator may be put to immediate use. All those accessories that will be supplied as part and parcel of the Simulator must be specified.
2.15 | Must be provided with a manual guide preferably in the English language.
2.16 | The unit must have a communication port / connectivity (Wired or Wireless), which will enable connection to a PC for automation test software and firmware upgrades
2.17 | The automation test software, used with the unit, must be able to integrate with various Computerised Maintenance Management System (CMMS) software’s, which is an advantage for future implementation
2.18 | Bidder to provide the name and make of various CMMS programs currently compatible with the unit