## Sheet1
Patient Simulator
1 | DESCRIPTION
1.1 | Patient simulator that can measure and simulate ECG Simulator, IBP Simulator, and Respiration parameters.
2 | TECHNICAL SPECIFICATION
2.1 | ECG Simulator
2.1.1 | Output at least 12 standard leads
2.1.2 | Amplitudes: 0.5, 1 and 2mV
2.1.3 | Output impedance: 500, 1000, and 2000Ω
2.1.4 | Adult or pediatric waveform
2.1.5 | Accuracy: ±1%
2.1.6 | Rate selectable: Between 30 and 300BPM
2.2 | IBP Simulation
2.2.1 | Output impedance: 350Ω
2.2.2 | Must be supplied with voltage level of ±10V from DC to 4000Hz
2.2.3 | Sensitivity: 5 to 40μV/V/mmHg
2.2.4 | Must be able to simulate a static pressure
2.2.5 | Different combinations of arterial and venous pressures must be possible
2.3 | Respiration
2.3.1 | Base line impedance: 500, 1000 and 2000Ω
2.3.2 | Rates: 15,30,40, 80, 120 and 0 BPM for apnea simulation
2.3.3 | Accuracy: Within 2%
2.3.4 | User selectable impedance variation
2.4 | General
2.4.1 | The unit must be portable, lightweight, handheld and have a protective carrying case
2.4.2 | The unit must be supplied with rechargeable batteries
2.4.3 | Display: At least 2 lines, 16 character LCD
2.4.4 | The Analyzer offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Analyzer may be put to immediate use. All those accessories that will be supplied as part and parcel of the Analyzer must be specified.
2.5 | Must be provided with a manual guide preferably in the English language.