## Sheet1
Ultrasound Leakage Tester
1 | DESCRIPTION
1.1 | Ultrasound leakage tester is designed to test the electrical safety of all types of diagnostic ultrasound transducers
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be able to test the electrical safety of all types of diagnostic ultrasound transducers totally independent of the ultrasound machines on which they are typically used.
2.2 | The unit offered must be able to Test the Upper & Lower Leakage Current Limits per Ultrasound Multi Function Relay (MFR) Transducer Specifications
2.3 | Operating Range 0.5 to 500 μA to Meet Transducer Specifications
2.4 | Large Graphic Display with Backlighting Allowing Easy Selection of Options & Setup of Parameters.
2.5 | User-Selectable Challenge (Test) Voltage (90 to 275 VAC) & Frequency (50 or 60 Hz).
2.6 | User-Selectable Test Limits
2.7 | Programmable Test Limits Based on Ultrasound Transducer Multi Functio Relay & Type
2.8 | ±1% F.S. Range Accuracy
2.9 | Auto Ranging for Enhanced Accuracy Over Entire Range
2.1 | Single Button Press for Full System Test
2.11 | Selectable Pass/Fail or Numerical Test Results
2.12 | User-Selectable Display Options
2.13 | Programmable Backlight Timer
2.14 | Flash Programmable, Field Upgradeable
2.15 | RS232 Interface
2.16 | PC Utility Software for Configuration Setup & Remote Control
2.17 | Test Results Printable With Optional Printer
2.18 | On-Board Clock & Calendar Function for Date/Time Stamp of Test Records
2.19 | On-Board Storage for atleast 100 Test Records.
2.1 | Audio-Visual Test Status Indication
2.11 | The tester must be able to operate on 220 Volt + 10% 50hz single phase a.c. power supply. The tester must be surge protected. The unit mains cable must be fitted with a SABS approved 15 amp plug top.
2.12 | The Tester offered must be supplied with a complete set of accessories as part of the unit, at no extra cost, in order that the Tester may be put to immediate use. All those accessories that will be supplied as part and parcel of the Tester must be specified. The unit must come in a carry case.
2.13 | Must be provided with a manual guide preferably in the English language.