## Sheet1
Dual Head Gamma Camera
Description
ITEM | Technical Specifications
1 | IT IS ALSO CRITICAL THAT THE PROSPECTIVE BIDDERS OF THIS SPECIALIZED RADIOLOGY EQUIPMENT MUST GUARANTEE THAT THE PRODUCT OFFERED HAS INTERNAL AND EXTERNAL NON – CORROSIVE MATERIAL/COMPONENTS, WHEN ARE EXPOSED TO EXTREME ENVIRONMENTAL FACTORS, SUCH AS THE SEA ETC
1.1 | The gamma camera shall be capable of:
Static, dynamic, whole body and gated planar acquisition and processing
SPECT (including gated SPECT) acquisition, reconstruction, processing and display.
2 | The gamma cameras will be required for, but not limited to, the following studies (i.e. 3phase, dynamic, static and whole body imaging for the various regions):
Head and Neck Studies
e) Thyroid Scans;
f) Parathyroid Imaging;
n) Cerebral perfusion and receptor imaging;
Chest and Cardiac Imaging
c) Myocardial perfusion imaging and thrombus imaging;
d) Gated cardiac blood pool imaging;
g) Lung ventilation and perfusion imaging including quantitative studies;
t) Scintimammography;
s) Sentinel Node imaging;
Abdominal Studies
i) Gastro-oesophageal reflux
j) Gastric emptying studies
k) Gastro-intestinal bleeding
b) Renograms and renal cortex functional imaging;
h) Milk scans;
p) Liver and spleen imaging
q) Hepatobiliary studies;
u) Adrenal Imaging;
Other Studies
a) Skeletal;
l) Gallium imaging for infection and tumours;
m) Infection / inflammation imaging with White Cell / leukoscan/ other available agents for this indication;
o) Tumour imaging in adults and children;
r) Lymphoscintigraphy
v) Pre- / post radionuclide therapy imaging
w) SPECT studies of any of the above.
3 | SPECT Camera
Must be upgradeable to hybrid SPECT/CT technology and this be priced as an option in the bid.
The CT component when installed should be capable of producing CT images that will be used primarily for attenuation correction and anatomic localization.
The supported CT technology options available should be clearly described, bidder to specify.
4 | Must be able to transfer images in DICOM and/or interfile format.
The systems should be able to connect to the planned PACS system (Department of Diagnostic Radiology) and be able to transfer images within the institution and other external institutions
5 | It shall also be possible to interface to the laser printer from the processing stations for hardcopies and to CD/DVD for soft copies
7 | The offered dual head SPECT gamma camera system shall consist of at least the following components:
a) Two large, rectangular field of view NaI detectors with a crystal thickness of 9.5mm
b) One set of low-energy high-resolution collimators
One set of low energy general purpose collimators
c) One set of medium-energy collimators
d) One set of high-energy general purpose collimators for imaging
e) One set of fan beam collimators
f) Pin- hole collimator
g) Storage and trolley for collimator change
h) Data acquisition computer and complete suite of acquisition software
i) Image Processing Computers and complete suite of standard processing software
j) Imaging table
k) ECG for cardiac imaging
l) Mammography pallet for scintimammography
m) Paediatric pallet
n) Integrated DVD/R RW
8 | The Gamma Camera shall have 2 rectangular NaI (TI) detectors with a useful field of view (UFOV) of at least 52 x 37 cm or greater.
9 | The number of photomultiplier tubes per detector shall be at least 56 or greater
10 | The intrinsic performance measured according to National Electrical Manufacturers Association (NEMA) specifications in normal operational mode for 140keV photons in the UFOV shall equal or better the following figures:
(a) Spatial Resolution of less than 4mm FWHM.
(b) Spatial linearity of 0.6 mm absolute and 0.2 mm differential.
(c) Energy resolution <10%.
(d) Uniformity of less then 4% integral and less then 3% differential.
11 | The extrinsic performance (measured according to NEMA specifications) for 140 keV with Tc-99m photons in the UFOV shall be quoted.
12 | The system spatial resolution with scatter (LEHR collimator distance 10 cm) for FWHM for CFOV should be < 9.0 mm and without scatter < 8.0mm.
13 | The system sensitivity per detector with Tc-99m LEHR collimator is > 160 cts/min/µCi.
14 | The reconstructed tomographic spatial resolution shall be, Tangential <9 mm; Radial <12 mm; Central <12mm.
15 | The energy range should be from 40 keV to 580 keV
16 | The variation of PHA response to a point source at 140 keV during a 360° rotation of the detector shall be 1% or better.
17 | The variation in count rate from a point source at 140 keV during a 360° rotation of the detector shall be 5% or better.
18 | The number of PHA windows which can be used to acquire and display data in a single image or image set must be specified with any limitation on dynamic, static, tomographic or gated tomographic acquisitions.
19 | Specify the number of discrete image sets which can be acquired from simultaneously applied PHA windows for static, dynamic, whole body, tomographic or gated tomographic acquisitions (“dual isotope acquisitions”).
Collimators:
20 | The following collimators must be included
a) A pair low energy (140 keV) high resolution parallel hole collimators.
One set of low energy general purpose collimators
b) A pair of medium energy (300 keV) general purpose parallel hole collimators
c) A pair of high energy (360 keV) general purpose parallel hole collimators
d) A pair of Fan beam collimators
e) Pin- hole collimator
21 | Collimators must be electronically coded in the acquisition system.
22 | The effective hole angulation of parallel hole collimators shall be preferably less than 0.15°.
23 | Storage rack and cart shall be supplied. The effective hole angulation of parallel hole collimators shall be preferably less than 0.15°.
Gantry and Detector Movements:
24 | All controls for movement of the detector head must be accessible for any head position.
25 | It must be possible for movements of the camera head and imaging table by a remote control device.
26 | The detector must remain in a fixed position unless activated by the operator.
27 | This must also apply in the event of a power failure.
28 | A patient must be able to be removed from the system in the event of a power failure easily
29 | The head must be equipped with appropriate sensors to detect contact of the head with the patient or an obstruction. This must prevent any further motion of the head.
30 | It must be possible to restart acquisition without loss of data.
31 | Independent detector angulations giving 90° to 180° for SPECT imaging shall be available. Tilting option shall also be available. Specify any other relative detector angles and any other limitations on system operation.
32 | The gantry and detector movements (e.g. rotational, translational) which are manual and those which are motor driven and remotely controlled shall be described. The speeds, range and directions of each of the motor driver movements shall be specified.
33 | Detector alignment with patient shall be automatically maintained while the detector / patient are varied.
34 | The system must have an auto contouring device for automatically sensing the patient-detector distance and adjusting the detector to maintain this distance constant as part of the basic system. The Bidder to state the technology employed.
35 | The acquisition modes for which the auto contouring modes can be used for:
a) SPECT;
b) Whole body.
c) Other specify
36 | Auto contouring device can be used with all the collimators
37 | It should be possible to deactivate auto contouring device
38 | State the distance between collimator and patient that the auto contouring device can achieve;
39 | The maximum and minimum radius of tomography shall be specified.
40 | Any additional general safety features must be separately described.
41 | For tomographic acquisition
a) Clockwise and anticlockwise motions shall be available.
b) Step and shoot plus continouos acquisitions shall be available.
c) Specify if other tomography angles are offered.
d) Specify the matrix sizes with which the system can acquire and store SPECT data and gated SPECT data.
e) An option to upgrade to SPECT /CT imaging SHOULD BE AVAILABLE and specified with regards to compatibility and technology
42 | Whole body Acquisition:
a) The range of scanning speeds should be between 5-60 cm/min.
b) Whole body scans are acquired continuous scanning motion
c) Maximum scan length shall be a minimum of 200 cm.
d) The matrix sizes available shall be a minimum 512 x 512
43 | Imaging Table:
a) The imaging table shall be suitable for static, dynamic, SPECT and whole body acquisitions.
b) The table shall be able to handle a patient load of at least 220 kg or more.
c) The imaging table should be able to be completely detached from the camera and wheeled to the side of the room.
d) The material used in the construction of the table; e.g. carbon fibre, aluminium. Specify the attenuation (%) of 140 keV photons.
44 | Physiological Triggering:
a) The system must be provided with a trigger input and appropriate software to enable acquired data to be gated for ECG gating cardiac studies.
b) Indicate if the ECG gating is controlled by an independent ECG monitor or an ECG monitor built into the gamma camera. Indicate whether it is possible to store the ECG signal as part of the patient file and if the ECG signal can be viewed afterwards
45 | Data Acquisition Hardware:
a) Specify the details of the computer supplied.
b) Indicate the hardware platform that the data acquisition workstation runs on (e.g. PC, MAC, SUN).
c) The latest computer hardware must be supplied.
d) Specify the CPU speed. The system must have a minimum speed of 2.8GHz.
e) Specify the amount of memory installed.
f) Specify the capacity of the hard disc drive.
g) Specify the operating system as well as the version that will be installed.
h) Specify any other storage devices available.
i) A flat panel LCD model with a minimum size of 19 inch must be supplied.
j) The monitor must have a minimum resolution of 1920 X 1200.
46 | Acquisition:
a) The system should be able to display images in persistence mode to enable ease of patient positioning before initiating acquisition and also be available for each detector head.
b) Rotate the acquiring image in multiples of 90° (i.e 90, 180, 270 and 360 degrees).
c) The range of acquisition software for the camera shall include at least the following: static, dynamic, multiple gated acquisition (planar) and gated SPECT, first pass, whole body, tomography.
d) The software shall allow the creation of predefined user protocols for routine acquisitions.
e) The following should be able to be specified in the pre-defined protocols: Study name, view name, energy window, matrix size, matrix depth, acquisition termination condition, image zoom, collimator, gantry and imaging table position.
f) It should be possible for images to be acquired from each head independently and stored separately.
g) It should be possible to store any study halted prematurely and to pause and restart acquisition without loss of data.
h) It should be possible for acquisitions to be terminable either by preset count, preset time, accepted heart beats, manually or by count density within a user-defined ROI.
i) Various matrix sizes and pixel depths should be available for the acquisition types that the system is capable of performing.
j) Various magnification factors should be available for zoomed acquisitions and indicate limitations thereof.
k) The system shall have a range of pre-defined radionuclides such as 99mTc, 67Ga, 57Co, 111In, 123I, 131I, 201TI and 81mKr and it shall be possible for the user to add radionuclides not included in the list.
l) The system shall be able to correct for isotope decay for tomographic as well as dynamic acquisition.
m) All acquisitions must be possible using conventional energy window settings that shall preferably be calibrated in keV.
n) For dynamic acquisitions indicate the minimum and maximum time per frame as well as the increments.
o) Indicate the number of different framing rates that would be possible to use during a single dynamic study.
p) Automatic body contouring shall be possible for whole body and SPECT images. Provide details of the method used for automatic body contouring.
q) Anterior and posterior images shall be acquired simultaneously for whole body acquisitions.
r) Gated planar and SPECT acquisition is required. A minimum of 16 frames is necessary. Specify the minimum and maximum number of frames that may be acquired during a cardiac cycle.
s) The user shall be able to define a window of acceptable beat lengths. Data from all cardiac cycles that fall outside this window shall be rejected.
t) The orbital movement for tomography shall include circular orbit, elliptical orbit and body contouring.
u) Gantry and bed positions shall be stored within the acquisition data file so that it can be used for patient re-positioning.
v) It should be possible for all collimators in tendered list to be used for SPECT.
w) Step and shoot acquisition shall be available and step angle increments and minimum and maximum shoot times and number of images indicated
47 | Quality Control Software:
The system shall have software for monitoring
a) the system as required by NEMA.
b) Software for the calibration of energy, linearity, uniformity and centre of rotation correction shall be available.
c) State if axial alignment correction is possible
48 | Image Display and ROI Management:
a) The system shall have several standard supplied colour tables, including gray scales, to display images.
b) The software should allow for the creation of user-defined colour tables.
c) It shall be possible to display several images simultaneously and to adjust display levels independently in each window.
d) It shall be possible to convert images from one matrix size to another for static, dynamic, tomographic and gated tomographic studies.
e) It shall be possible to create a new dynamic study from frames of an existing dynamic study.
f) Summation of contiguous frames in a dynamic series shall be executable.
g) It must be possible to produce condensed images from a dynamic series.
h) Indicate if variable width profile generation and analysis thereof is possible.
i) Spatial and temporal smoothing of images shall be possible.
j) Image algebra on static and dynamic images shall be available.
k) Image rotation and mirroring shall be available as well as linear translation for static, dynamic and tomography acquisitions so that correction for patient motion is possible.
l) It shall be possible to view SPECT data in a 3-D volume/cine display.
m) The system shall have software for the creation of ROI’s on images or summed images using manual methods, preset geometrical shapes, isocount contouring and image thresholding.
n) Indicate if it is possible to store ROI’s in the database so that they can be retrieved and manipulated at a later stage.
o) Indicate if ROI’s can be superimposed on a cine display of images and any limitations thereof.
p) Indicate if two separate cine-gated cardiac displays (e.g. stress and rest) can be synchronized so that corresponding frames are displayed simultaneously.
q) ROI arithmetic shall be performable and data production on total counts, size and average counts per pixel shall be possible. ROI manipulation and mirroring shall also be possible
r) Generation of activity-time curves from ROI’s shall be possible. The storage and display of such curves shall also be possible.
s) Image quantitation ability shall be available
t) It shall be possible to perform curve algebra, to smooth, integrate and differentiate curves.
u) Indicate if curve fitting is possible and if so, indicate the types.
49 | Reconstruction Software:
a) Reconstruction of tomography images shall be possible using filtered back projection as well as iterative techniques. Iterative reconstruction software shall be possible
b) Indicate the maximum tolerable centre of rotation offset that can be corrected by software.
c) Indicate if scatter corrections are available and what this entails.
d) Specify the method(s) of attenuation correction that are available.
e) SPECT reconstruction and storage of axial, coronal, sagital and oblique planes shall be possible.
f) SPECT reconstruction of circular, elliptical and body contour orbits shall be possible.
g) Indicate the filters that are available for reconstruction purposes.
h) Indicate if alignment of sequential tomographic images is possible using either landmarks or manually.
I) CT reconstruction and storage of axial, coronal, sagittal and oblique planes shall be possible.
50 | Clinical Processing Software:
Indicate if the system has the capability to perform the following processing functions for dynamic renal imaging:
a) Calculation of relative function;
b) Calculation of retention function;
c) Calculation of parenchymal transit time;
d) Calculation of output efficiency;
e) Calculation of whole kidney transit time;
f) Determination of furosemide response;
g) Determination of the washout halftime;
h) Generation of Patlak-Rutland plots and
i) Determination of the perfusion index.
51 | Indicate if the system has software for calculation of ventilation/perfusion ratio in lung imaging and calculation of washout rates from ventilation studies
52 | Indicate if the system is capable of
a) Esophageal parametric images and the calculation of the transit time.
b) Gastric emptying rate and halftime can be calculated including use of geometric mean images.
c) Parathyroid subtraction imaging software is available.
d) Brain Processing Software.
53 | For cardiac
a) Create Bull’s-eye images for myocardial perfusion
b) 4DM/QGS cardiac processing software.
c) Software for the 3-axis display of myocardial perfusion images.
d) MUGA can be determined both semi & fully automatically and if the Fourier phase & amplitude images can be displayed.
e) Beat histogram can be displayed.
f) MUGA calculations can be performed for LVEF, PFR, PER and regional EF.
g) Calculate First pass RVEF.
h) Perform First pass shunt quantitation.
54 | It shall be indicated if software updates will be supplied and installed and what the relevant costs will be.
55 | File Handling and Data Processing:
a) The system must have standard directory and file handling utilities.
b) The system must have archiving and indexing software and allow searches using several different criteria such as patient name, ID number, study type etc.
c) Archiving can be initiated e.g. automatically or manually.
d) The system shall have the software to convert images to TIFF, JPEG, GIF or AVI formats that are compatible with MS Windows.
e) The system shall have the capability and necessary software for CD or DVD writing.
f) Installation of third party software such as antivirus software shall be possible.
g) The system shall be capable of connecting to other systems using standard Ethernet protocols.
h) The system shall be able to reliably import and export data from and to other processors currently in use with Interfile or DICOM facilities.
i) The system shall have the necessary TCP/IP software to make communication by a standard network system possible.
j) Indicate if automatic transfer of completed acquisitions is possible.
k) Indicate if it is possible for the user to transfer studies from the system to a remote station (push) and/or remotely transfer studies from the system (pull).
l) The Bidder to provide details about the image archiving and storage facilities, as well as the number of 512 x 512 resolution images that can be stored, on the internal hard drive(s). The hard drive(s) shall be large-capacity industry standard, and easily upgradeable.
m) It shall also be possible to store/archive images on commercially available CD-ROM or DVD disks in a generally recognizable and readable DICOM format.
56 | DICOM COMPATIBLE LASER CAMERA.
DICOM Compatible networked laser camera must be provided. (455 DPI or better, for printing on 18 x 24cm film size is required).
51 | APPLICABLE DOCUMENTS
The offered product shall comply with the National Year 2000 Conformity Requirements as defined in the Republic of South Africa. PD2000-1: 1998 Year 2000 definition. Please clearly state if the offered product complies
52 | POWER SUPPLY
The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
A 3 phase in / out power supply must be offered. State details.
The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
The lifetime of UPS must be at least 15 years.
An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
The mains cable of the unit tendered for must be SABS colour coded. State details.
The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
Maintenance free batteries must be offered.
58 | OPTIONS
The Bidder is to offer the following as optional extras (state price and details) for each separately. The offered options shall be fully functional options that require no additional hardware or software to function:
OPTION A
Compatible CT component technology to upgrade the currect SPECT to hybrid SPECT/CT
The different available options of CT technology packages that can be incorporated on the system to upgrade to a hybrid SPECT/CT camera should be clearly stated and described for comprehensive consideration
OPTION B
Image Registration Software
It must be able to automatically register SPECT and CT images
It must be able to register images which are obtained from other modalities such as MRI, CT and PET and other gamma cameras.
OPTION C
If there are any additional available software or hardware options not specified in this bid, the bidder is requested to motivate the benefits of each option.
OPTION D
If a separate persistence scope is not supplied with the main offer, quote the price for a persistence scope.
OPTION E
Must be supplied with fully automated collimator exchange system.