## Sheet1
70cm Wide Bore Magnetic Resonance Imaging(MRI) Unit (1.5T) Tesla
Item | Technical Specifications
1 | MAGNET
1.1 | The Magnet Field Strength shall be at least 1.5 Tesla active shielded
1.1.1 | State available upgrade path for unit offered (without magnet replacement). Please also state if the magnet on offer is the latest Technology and Release
1.2 | The Magnet shall be short bore cylindrical type and design
1.3 | Bore length including all covers shall not exceed 180 cm
1.4 | State internal bore dimensions (L x W x H)
1.4.1 | Maximum Vision Angle
1.5 | Magnet homogeneity shall meet the following specifications using the standard deviation VRMS (volume root-meansquare) using 24 plane plot measurement method:
1.5.1 | 50 cm DSV state
1.5.2 | 48 cm DSV state
1.5.3 | 45 cm DSV state
1.5.4 | 40 cm DSV state
1.5.5 | 30 cm DSV state
1.5.6 | 20 cm DSV state
1.5.7 | 10 cm DSV state
1.6 | Magnetic field stability shall be less than 0.1 ppm/hour or better
1.7 | The 5 Gauss/0.5mT fringe field shall be contained in an area of typically 2.5 m (radial) by 4.0m (axial). State actual area
1.8 | State helium boil-off details
1.8.1 | State typical cryogen refilling interval
1.8.2 | State maximum helium capacity
1.8.3 | State the weight of magnet with cryogens, table and covers (kg) in full operation
1.8.4 | The bore diameter shall be 70cm measured at the center in an operational mode
1.9 | The tenderer shall guarantee and ensure that there is no magnet vibration of the system, on the current location, with all possible imaging sequences. Give details.
1.9.1 | Space required (Magnet, Electronics and Control room)
2 | MAGNET SAFETY
2.1 | Magnet shall be equipped with quench exhaust leading to the outside of the building in event of Magnet quench to prevent injury to staff and patient
2.2 | Magnet shall be equipped with emergency ramp down unit for fast Magnetic Field reduction.
2.3 | Any other safety regulations by Environmental Safety Law and Regulations must be applied
2.4 | Successful bidder must provide a MRI user license obtained from Radiation Directorate
2.5 | Metal detector to be build into the magnet entrance door.
2.6 | Examination room shall be marked off by warning labels/lights
3 | GRADIENT SYSTEM
3.1 | Actively shielded hi-performance non-resonant gradient coil system
3.2 | Gradient amplitude/peak strength shall be at least 30mT/m measured per real axis plateau (100% duty cycle) not effective gradients
3.3 | Gradient duty cycle shall be 100%
3.4 | Minimum rise time from 0 to max. gradient amplitude shall be 240 micro-seconds or better
3.5 | Gradient slew rate shall be 120 T/m/s or better measured per real axis plateau, not effective values
3.5.1 | Shortest TE (SSh DWI, b=1000, 128 matrix)
3.5.2 | Shortest TE / TR (2D, 128 matrix) : 2D GRE
3.5.3 | Shortest TE / TR (2D, 128 matrix) : 3D GRE
3.5.4 | Shortest TE / TR (2D, 128 matrix) : 2D GRE EPI
3.5.5 | Shortest TE / TR (2D, 256 matrix) : 2D GRE
3.5.6 | Shortest TE / TR (2D, 256 matrix) : 3D GRE
3.5.7 | Shortest TE / TR (2D, 256 matrix) : 2D GRE EPI
3.5.8 | Maximum scan matrix
3.5.9 | Maximum recon matrix
3.5.10 | Maximum scan matrix
3.5.11 | Maximum recon matrix
3.5.12 | Increment steps for matrix
3.5.13 | Maximum number of slices
3.5.14 | Minimum slice thickness (3D)
3.5.15 | Duty cycle
3.5.16 | Gradient linearity at 30 cm DSV
3.5.17 | Gradient linearity at 40 cm DSV
3.5.18 | Gradient linearity at 50 cm DSV
3.6 | Gradient upgrades to different levels should be possible with out replacing magnet.
3.6.1 | State high-end gradient specifications for optional gradient upgrade:
3.6.2 | Max slew rate and amplitude and FOV at max slew rate and amplitude
3.6.3 | Cooling for gradient coil; state type of coolant
3.7 | The output linearity of the gradient amplifiers should be no worse than + 0.1% of peak
3.8 | A Child Friendly scanning enviroment is required. Noise should be kept to a minimum. State advanced ergonomic features that characterize the configuration offered.
3.9 | Free choice of flip angle while maintaining signal to noise ratio to be supported. Specify.
4 | RADIOFREQUENCY SYSTEM
4.1 | Resonance frequency shall be 63.86MHz (1.5 Tesla)
4.1.1 | Direct Radio Frequency (RF) Transmitter System: Digital signal generation and processing.
4.1.2 | RF system shall be digital transmitter and receiver design
4.1.3 | Maximum power output of transmitter amplifier rating shall be 28kW
4.1.4 | State the bandwidth of the RF transmitter
4.1.5 | The system shall be equipped with RF fault protection limiting RF output in event of malfunction
4.1.6 | Frequency resolution of the RF synthesizer shall be 0.35 Hz or better
4.2 | The receive components shall be integrated into the magnet housing
4.2.1 | The system should have at least 18 independent RF receiver channels or better
4.2.2 | Phase resolution shall be 0.1 degree/bit or better
4.2.4 | Noise of the preamplifier to be ≤ 0.5 decibels
4.2.5 | State maximum receiver bandwidth of each receiver channel
4.3 | Can all of these coil elements be seamlessly used in one study without coil changes and without repositioning of the patient
4.4 | State maximum number of simultaneously connected coil elements. Preference will be given to highest number. Indicate coil combination.
4.5 | State maximum receiver bandwidth of each receiver channel.
4.6 | Sampling rate of each ADC to be about 80 MHz
4.7 | Receiver signal resolution should be 32 bit
5 | RF COILS
5.1 | The bidder shall supply the latest Integrated Coil Technology (i.e. GE = GEM Technology, Toshiba = Atlas Technology, Siemens = TIM Technology and Philips = dStream Technology). Please state how the Coil Philosophy of the unit contributes and improve the Image Quality and Workflow .
5.2 | Coil pre-amplifiers shall be on the patient table connector and coils shall be interchangeable and be of light construction.
5.3 | Diameter of 70cm is required (larger bore will be preffered because of the patient profile at this institution e.g. obese, claustrophic, paediatrics etc.) NB: The bore to be measured at centre of the system and not on the flare!
5.4 | Give standard surface coils supplied with system
5.4.1 | Maximum coils connection to do full body scan. Give details of coils and connections (short cable coil connections)
5.5 | Connected coils must be detected automatically
5.6 | It must be possible to select active coil or elements from the main console
5.7 | Special coils for the following applications should be available with the system (state price separately). State standard coils with the system.
5.7.1 | Body (Chest, abdomen, pelvis)
5.7.2 | TMJ kit
5.7.3 | Head
5.7.4 | Head-neck (without repositioning patient)
5.7.5 | C/T/L spine ( State what coil is used for imaging of the total spine cervical, thoracic and lumbar )
5.7.6 | Whole spine
5.7.7 | Neck soft tissue
5.7.8 | Extremities and joints: wrist, hips, shoulder, knee, foot/ankle,
5.7.9 | Peripheral array coil to be included in the price
5.7.10 | Prostate, colon, cervix (Endorectal)
5.7.11 | Whole body vascular (Peripheral vascular array)
5.7.12 | Peripheral MR Angiography
5.7.13 | Neurovascular
5.7.14 | Cardiac
5.7.15 | All paediatric applications
5.7.16 | Supply a cupboard in the room for storage of all coils provided or alternatively sufficient coil cabinets to house the coils.
5.7.17 | State all other SW packages that comes standard with the system.
5.7.18 | Other features
Coils to facilitate optimal facial, calvarial and intracranial imaging in the child from 1.5-60kgs, including visualisation of the orbits, optic nerves, optic tracts, pituitary fossa, hypothalamus, temporal lobe, posterior fossa and cerebral cortex. Coils to provide optimal imaging including, but not limited to the following clinical settings:
6 | PATIENT SUPPORT/TABLE and MANAGEMENT/ PATIENT COMFORT
6.1 | 2 X Patient table shall be dockable
6.2 | Vertical table movement
6.2.1.1 | Minimum table height from floor
6.2.1.2 | Maximum table height from floor
6.2.1.3 | Maximum vertical speed
6.2.2.1 | Horizontal table movement
6.2.2.2 | Maximum horizontal range of table movement
6.2.2.3 | Maximum speed:
6.2.2.4 | Accuracy of repositioning
6.3 | Dual table control panels shall be located at either side of aperture/gantry for easy access. State if controls are available on the rear of the magnet
6.4 | Three alignment light beams for anatomical references in axial, coronal and sagittal planes
6.5 | Table movement shall be controlled from both gantry and operator console
6.6 | Patient table shall be equipped with manual override for quick removal of patient from the magnet-bore in case of emergency
6.7 | Maximum load shall be 200kg or better. State actual maximum load.
6.8 | Physiologic measurement unit essential with display of ECG, respiration and pulse at the main console
6.9 | Two way in-bore intercom system shall allow communication with patient while gradient is running
6.10 | In-bore music/stereo system shall be included
6.11 | Hand held alarm button for patient signalling
6.12 | Table restraining/immobilization straps are required
6.13 | Separate Patient transport stretcher –MR compatible Max load 200kg to be supplied.
6.14 | Integrated music for patient, including CD interface to be included
6.15 | Variable patient lighting to be included.
6.16 | Child friendly interior decoration to be included.
6.17 | State of the art technology is required to allow young patients to see outside the magnet during the examination. Give details
6.18 | MR compatible sand bags
6.19 | Two comfortable patient mattresses to be included
7 | ADDITIONAL WORKSTATION: DOCTORS CONSOLE
7.1 | In addition to the acquisition console, additional, 2 x dedicated Workstations (Doctors Workstation) will be supplied (or better solution).
7.2 | The dedicated workstation will be used for all post processing. The system must included licenses for all 3D reconstruction and post processing requirements. Please give details and options where applicable
7.3 | Please state if the additional workstation has the same user interface as the console
7.4 | Please state if the additional workstation can have the same software packages as the console
7.5 | High quality read/write CD/DVD storage device required
8 | OPERATOR USER INTERFACE
8.1 | 23 inch hi-resolution colour LCD flat-panel flicker-free monitor with undistorted image display required
8.2 | Monitor 1024 x 1024 dot resolution required
8.3 | User interface providing flexible multi-tasking in foreground or background (scanning, filming, reconstruction)
8.4 | All standard image processing feature are required in addition the following advance features must be available:
8.4.1 | Multi-planner reformatting (MPR)
8.4.2 | Multi-projection volume rendering (MPVR)
8.4.3 | Maximum intensity projection (MIP)
8.4.4 | MR Angiography processing
8.4.5 | 3D surface rendering
8.4.6 | MR Hydrography MRCP/Urography/Myelography
8.4.7 | Image add/subtract
8.4.8 | CT image display
8.4.9 | CT image integrating must be offered.
8.4.10 | State any special features to improve productivity and consistency (e.g. automated scanning procedures)
9 | ACQUISITION PARAMATERS, APPLICATION SOFTWARE AND IMAGING TECHNIQUES
9.1 | State slice and slab thickness:
9.1.1 | 2D
9.1.2 | 3D
9.2 | State number of slices:
9.2.1 | 2D
9.2.2 | 3D
9.3 | Variable Field Of View required to maximum of 48cm or better. State
9.4 | STANDARD/CONVENTIONAL AND FAST IMAGING TECHNIQUES/ SEQUENCES REQUIRED FOR ALL THE FOLLOWING APPLICATIONS WITH DEDICATED POST PROCESSING
a | State standard sequences in offered software package
b | State maximum parallel imaging factor possible for offered software package (and compatible coils)
c | State sequences that are compatible with parallel imaging
d | State optional software packages. State details and costs
9.4.1 | Neuro:
i | Brain
ii | Conventional imaging sequences
iii | Spectroscopy (Single voxel and multi voxel). State if compatible with parallel imaging 2D / 3D
iv | Diffusion weighted
v | Perfusion weighted. Sate sequences used.
vi | Fast MRI sequences. Sate sequences used.
vii | Tractography / Fibre tracking
viii | Spine
9.4.2 | Cardiac:
i | Conventional sequences
9.4.3 | Abdomen:
i | Conventional sequences
ii | Virtual endoscopy
iii | MRCP
iv | Small bowel studies
v | Dynamic liver contrast studies
9.4.4 | Chest:
i | Conventional sequences
ii | Mammography
iii | Dynamic (3D FSGRE etc.)
iv | Biopsy capability
v | Pulmonary ventilation studies
9.4.5 | Pelvis:
i | Conventional sequences
ii | Prostate imaging (including spectroscopy)
iii | Dynamic pelvic floor imaging
9.4.6 | Vascular:
i | TOF (2D/3D)
ii | PC (2D/3D)
iii | Contrast enhancement for:
iv | Neck vessel studies
v | Intracranial vessel studies
vi | Aortic arch and branches
vii | Abdominal aorta and outflow
viii | Upper limbs
ix | Pulmonary arteries
x | Renal arteries
xi | Multistep MR angiography
xii | Lower limps
9.4.7 | Head and neck:
i | Conventional sequences
ii | IAM's
iii | TMJ's
10 | PERFORMANCE EVALUATION PHANTOMS REQUIRED / QUALITY CONTROL
10.1 | Performance evaluation software and appropriate phantoms required for evaluation of image quality required as per licensing conditions
11 | SAFETY
11.1 | The Magnet system shall include an emergency ramp down unit (RDU) for fast reduction of The Magnetic field.
11.2 | The Magnet shall have quench bands that contain The fringe Field in The event of a Magnet quench.
11.3 | Real-time SAR calculation shall be performed by software to ensure that RF power levels comply with regulatory guidelines and be displayed on each image.
11.4 | The system shall have manual override of the motor drive for quick removal of the patient from the magnet bore.
11.5 | The system shall have RF fault protection hardware (RF transmit enable limit and RF power and duty cycle limit) to limit The RF output in The event of equipment malfunction.
11.6 | Walk through metal detector
12 | ADDITIONAL EQUIPMENT
12.1 | MRI compatible ECG cables required
12.3 | MRI integrated pulse oximeter for adults and children
13 | MR FLUROSCOPY
13.1 | a MR Fluoroscopy mode with real-time acquisition and display techniques is required as an option. Please State the soft and hardware required if this is optional.
14 | STANDARD ACCESSORIES
List of standard accessories included but not specified in the tender.
14.1 | Incubator for Neonatal Imaging
14.2 | MRI compatible contrast injector interphased with MR unit
14.3 | 200 syringes and 200 matching syringe lines
15 | TECHNOLOGY
15.1 | Innovation. Give explanation
15.2 | Unique. Give explanation
15.3 | Clinical benefits. Give detailed explanation
16 | OPTIONAL ACCESSORIES
16.1 | The tenderer to give full description and pricing of optional accessories available for the system. Each item must be priced separately
17 | MRI COMPATIBLE ANAESTHETIC UNIT
17.1 | MRI Compatible Anaesthetic unit for General Anaesthetics of the paedeatric patients is required. Please state the specifications of the unit on offer. Please also note the warranty and customer support agreement terms.
17 | LASER CAMERA
17.1 | A state of the art dry laser camera to be included. Specify: Make Model
17.2 | UPS for the laser camera to be included
B General Issues
1 | SAFETY STANDARS OF THE SYSTEM
1.1 | MR Unit should be CE marked
1.2 | MR Unit should be manufacture according to FDA 1020.30 regulations
1.3 | MR Unit must Comply to IEC 601
1.4 | The equipment/system must be approved and licensed by Radiation Control.
1.5 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
1.6 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
2 | IT INTEGRATION
The system must have the following DICOM 3 compatibility:
2.1 | DICOM Send / Receive
2.2 | DICOM Storage Commitment (SC):
2.3 | DICOM query/retrieve
2.4 | DICOM worklist
2.5 | DICOM Basic Print
2.6 | DICOM MPPS (Modality Performed Procedure Steps)
2.7 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
2.7.1 | Scheduled Workflow (SWF)
2.7.2 | Patient Information Reconciliation (PIR)
2.7.3 | Simple Image and Numeric Reports (SINR)
2.8 | System to be integrated to the existing PACS
3 | AFTER SALES CUSTOMER CARE AND SERVICE
3.1 | State capacity and capability of your after sales service
3.2 | Supply full installed base of same or similar systems on offer
3.3 | A full Remote Services funtionality to reduce downtime. Give details
3.4 | Spares turn-around to achieve 98% Up-time. Give details
3.5 | The up-time of the unit must be better than 98%, excluding scheduled preventative maintenance and software upgrades, measured on a quarterly basis.The percentage lower than 98% will be added to the warranty period. A sliding scale penalty clause will form part of the service contract. This will result in the maintenance payment being reduced by a pro rata amount that the up-time is less than 98%.
3.6 | Up-Time is defined as follows: 24/7; i.e. 365days times 24 hours = 8760 Hours. A down time of 2% relates to 175 hours per annum.
3.7 | No part shall be second hand or refurbished.
3.8 | Spare parts must be guaranteed available for the specified life of the equipment, with a minimum of ten years.
3.9 | It must be guaranteed that no additional equipment, parts or software, excluding consumables, is required to operate the equipment specified in this tender.
4 | AIR CONDITIONING
4.1 | State power
4.2 | Specialised air conditioning will not be supplied by the hospital to the technical room and air conditioning will not be supplied to the examination room by the hospital. Vendors must include any additional air conditioning required in their offer. This equipment will remain the responsibility of the vendor interim of guarantees and service as described in item 5 below. State cooling required.
4.3 | The Chiller unit needed to support the cryogenic system must been seen as part of the system in term of supply, maintenance, support and guarantees.
5 | POWER REQUIREMENTS
5.1 | Tenderers must state the power requirements for the system of offer.
5.2 | Please provide the price for a UPS (uninterrupted power supply) to the unit allowing full operational functionality. State the size of the UPS and specification required for installation (Optional Item). The computer systems should be standard on a UPS
6 | SERVICE AND MAINTENANCE
6.1 | The thirty six (36) months guarantee must be included in the unit price of the equipment. The purchase pricing schedule must be completed in full.
6.1.1 | All equipment, materials and workmanship provided under this contract shall be unconditionally guaranteed for a minimum period of thirty six months (36) months (including the Helium ) from the date of commissioning. State guarantee period. 98% Up-time must be guaranteed. All costs associated with any preventative and safety inspection required during the guarantee period or any quality assurance test required during the guarantee period shall be included in the contract.
6.2 | Specify the total cumulative cost for the 5 year all inclusive service contract.The maintenance contract pricing schedule must also be completed in full.
6.2.1 | A fully comprehensive 5 year, service and maintenance contract, including all labour and parts(including Helium) for years 4-8 years must be Quote for separately. 98% Up-time must be guaranteed.
7 | TRAINING
7.1 | Applications specialist should train all users for 2 weeks on site during commissioning. During the 12 months following commissioning, session of 4 days be scheduled for follow-up training.
7.2 | On-going application training by Applications Specialist of 4 days a year must be provided at no additional cost.
7.3 | Full details to be given of all training specified and offered
8 | ROOM AND SITE PREPARATION
8.1 | A Site visit is required to determine that proposed site is suitable
8.2 | Building alterations are to be indicated on a layout drawing with a full scope of work. Quotation to be supplied separately
8.3 | All costs associated with the provision of such information in whatever form required including the necessary systems planning expertise shall be for the cost of the successful Tenderer
8.4 | Tenderers are required to submit a detailed plan indicating layout of proposed equipment configuration
9 | UPGRADES
9.1 | Free upgrades of hardware provided in the guarantee period