## Sheet1
PET CT (16 slices) Stand Alone
Item | Technical Specifications
1 | Gantry and System
1.1 | Height between 160 cm to 210cm
1.2 | Width between 190 cm to 240cm
1.3 | Depth between 100 cm to 150 cm
1.4 | Patient port atleast 70 cm
1.5 | Total system weight must be from 3 kg to 5 kg
2 | Patient Handling System
2.1 | Minimum patient load of 200 kg
2.2 | Max co-scan range of atleast 180 cm and above
2.3 | Length between 350 cm to 390 cm
2.4 | Width between 30 cm to 60 cm
3 | PET Acquisition
3.1 | PET Service Tools
3.2 | 3D Acquisition
3.3 | Online histograming (static beds)
3.4 | Automated multi-bed acquisitions
3.5 | Acquisition with offline histogramming supports dynamic and gated images
4 | PET Reconstruction
4.1 | Asynchronous processing/reconstruction
4.1 | 3D image-based scatter correction
4.1 | Metal artifact reduction for cardiac imaging
4.1 | 3D iterative reconstruction
4.1 | ToF
4.1 | FORE + 2D Backprojection
4.1 | Rb-82 Prompts Gamma Correction
4.1 | Reconstruction of gated datasets
4.1 | Reconstruction of dynamic datasets
5 | PET Detector Assembly
5.1 | Detector material must be LSO
5.2 | Detector element dimension of a minimum 4 x 4 x 20 mm
5.3 | A minimum of 160 detector elements per block
5.4 | 4 Photomultiplier tubes (PMTs) per block
5.5 | Detector ring diameter between 820 to 860 mm
5.6 | 48 Blocks per detector ring
5.7 | Transaxial FoV of atleast 70 cm
5.8 | 2 mm Plane spacing
5.9 | Atleast 620 and above detectors elements per ring
5.10. | Atleast 35 and above detector element rings
5.11 | 144 total detector blocks
5.12 | Total number of detector elements minimum of 24 336 (4-Ring TrueV
5.13 | Axial FoV minimum of 160 (4-Ring TrueV)
5.14 | Atleast 75 and above image planes
6 | PET NEMA 2012 Performance
6.1 | Scatter fraction (peak NECR) a minimum 35%
6.2 | Scatter faction (low activity conc.) a minimum 30%
6.3 | Count rate accuracy (Δr mean bias ≤22 kBq/cc) a minimum ±4%
6.4 | Count rate accuracy (Δr mean bias @ peak NEC) a minimum ±4%
7 | PET Data Acquisition/Processing
7.1 | Coincidence window of atleast 3.8 nsec
7.2 | System energy resolution of ≤12% FWHM
7.3 | System time resolution (ToF) of 500 or better psec typical
7.4 | Acquisition modes (static multi-bed, dynamic FlowMotion of up to 20 and above pases and variable speed per bed region)
8 | PET Reconstruction Times
8.1 | Reconstruction time per bed (200 x 200)
8.2 | ToF 30 sec
8.3 | UltraHD•PET 35 sec
9 | Spatial Resolution–Axial
9.1 | Standard Processing (256x256)
9.2 | FWHM @ 1 cm (mm) less than 5.5 mm
9.3 | FWHM @ 10 cm (mm) minimum 6 mm
9.4 | HI-REZProcessing (400x400)
9.5 | FWHM @ 1 cm (mm) less than 5 mm
9.6 | FWHM @ 10 cm (mm) minimum 6 mm
10 | Spatial Resolution–Transverse
10.1 | Standard Processing (256x256)
10.2 | FWHM @ 1 cm (mm) less than 5.5 mm
10.3 | FWHM @ 10 cm (mm) minimum 6 mm
10.4 | HI-REZProcessing (400x400)
10.5 | FWHM @ 1 cm (mm) less than 5 mm
10.6 | FWHM @ 10 cm (mm) minimum 5 mm
10.7 | Sensitivity (cps/kBq) less than 5
10.8 | Peak NEC rate (kcps) 107 @ ≤30 kBq/cc
10.9 | Peak trues rate (kcps) 380 @ ≤46 kBq/cc
11 | Software
11.1 | Time of Flight (ToF)
11.2 | CT extended FoV up to 80 cm
11.3 | PET offline reconstruction
11.4 | Easy and intuitive scan parameter setting (FAST Scan Assistant)
11.5 | Direct scan parameter adjustment (FAST Adjust)
11.6 | Patient Dose management (CT)-(CARE kV; CARE Profile;CARE Dashboard;CARE Child; CARE Filter; CARE Topo; CARE Dose4D™; CARE Bolus CT)
11.7 | Workstream 4D (3D-Recon)- 4D workflow with direct generation of axial, sagittal, coronal, or double-oblique images from standard scanning protocols
11.8 | Visualization, volumetric analysis, and registration/fusion of PET·CT studies (syngo TrueD Basic)
11.9 | Volume Rendering Technique (syngo VRT)
11.10. | Volume Calculation (syngo Volume Calculation)
11.11 | Surface Shaded Display-syngo 3D SSD (Surface Shaded Display)
11.12 | Archiving and Network (syngo Archiving and Network)
11.13 | Dynamic Evaluation (syngo Dynamic Evaluation)
11.14 | DICOM viewer–Included on each CD; automatically started on the viewer’s PC (DICOM viewer–Included on each CD; automatically started on the viewer’s PC)
12 | CT General Specifications
12.1 | Aperture 70 cm and above
12.2 | Scan field (HD FoV 78 cm (extended FoV) / Diagnostic FoV min 50 cm)
12.3 | Rotation times 1 sec or better
12.4 | 16 slices or more per rotation
12.5 | Temporal resolution-down to 83 ms
12.6 | Multislice Detector
12.7 | Minimum of 24 detector rows
12.8 | Minimum of 20 detector electronic channels (DAS)
12.9 | Minimum of 14 500 detector elements
12.10. | 1 400 Total channels/slice
12.11 | Number of projections up to 2320/360°
12.12 | Generator maximum power minimum 50kW-80 kW
12.13 | X-ray tube
Tube current 20-666 mA
Tube voltages 70, 80, 100, 120, 140 kV
Tube focal spot 0.7 x 0.7 mm/7°, 0.9 x 1.1 mm/7°
13 | CT Reconstruction System
13.1 | Reconstruction speed minimum of 16-20 images/s
14 | CT Topogram
14.1 | Length of topogram field minimum of 128-2,200 mm
14.2 | Topogram scan times of 1.5-20 sec
15 | CT Sequence Acquisition
15.1 | Reconstructed slice widths in sequence acquisition (0.6, 0.75, 1, 1.2, 1.5, 2, 2.4, 3, 4, 4.8, 5, 6, 7, 7.2, 8, 9, 10,12, 14.4, 15 mm)
15.2 | Scan times (360° full scan) 1 sec or better
15.3 | Scan times (260° partial scan) 0.24, 0.36, 0.72 sec
15.4 | Atleast 90 uninterrupted scans per range
15.5 | Atleast 30 ranges per protocol
15.6 | Scan cycle time 0.75-60 ses (±10%)
16 | CT Spiral Acquisition
16.1 | Reconstructed slice widths in spiral acquisition (0.6, 0.75, 1, 1.5, 2, 3, 4, 5, 6, 7, 8, 10 mm)
16.2 | Scan times (360° full scan) 5 sec or better
16.3 | Slice increment 0.1-10 mm
16.4 | Pitch factor 0.35 -1.5, 0.15, 0.09
16.5 | Spiral scan time max. atleast 100 sec or better
16.6 | Minimum of 30 ranges per protocol
17 | CT Reconstruction
17.1 | Real-time display supported during spiral acquisition
17.2 | Reconstruction field 5-50 cm, 5-78 cm (HD FoV)
17.3 | Reconstruction matrix 512 x 512
17.4 | HU scale 1024 to 3071
17.5 | Extended HU scale 1024 to 3071
17.6 | Reconstruction time min 16-20 images/s
18 | CT Image Quality
18.1 | Low-contrast resolution (CATPHAN 20 cm) / Technique–10 mm, 120 kV
18.2 | Sequence
18.2.1 | Object size 5 mm
18.2.2 | Contrast difference 3 HU and above
18.2.3 | Dose at surface 10.7 mGy@180 mAs
18.3 | Spiral
18.3.1 | Object size 5 mm
18.3.2 | Contrast difference 3 HU and above
18.3.3 | Dose at surface 13.1 mGy@180 mAs
18.4 | High-contrast Resolution
18.4.1 | Technique (160 mA,/120 kV,/0.5 s,/2.4 mm/(12x1.2mm)/UHR/Kernel H70h)
18.4.2 | Plane 0% MTF (±10%) 17.4 lp/cm
18.4.3 | Plane 2% MTF (±10%) 16.4 lp/cm
18.5 | Scanner Room Environment
18.5.1 | Temperature range 20-26° C (68-79°F)
18.5.2 | Relative air humidity without condensation (recommended level 30-70%)
18.5.3 | HVAC requirement (heat dissipation during operation) 2.0 kW (6,667 BTU/hr)
18.5.4 | Air pressure 750-1,060 mbar
18.5.5 | Variation in floor levelness not to exceed 12.7 mm (0.5 in) over 5.2 m (17 ft)
18.5.6 | Minimum room size 4.6 m x 7.3 m (15 ft x 24 ft)–(requires angling gantry and PHS)
18.5.7 | Recommended room size ≥5.0 m (16.5 ft) x ≥7.9 m (26 ft)
19 | Utility Room Environment
19.1 | Temperature range 18-30° C (64-86° F)
19.2 | Relative air humidity without condensation (recommended level 30-70%)
19.3 | HVAC requirement removal of 5.0 kW (17,100 BTU/hr)
19.4 | Gantry heat exchanger–max. 15 kW to water cooling environment
19.5 | Minimum room size 3.2 m (10.5 ft) x 2.1 m (7 ft)
20 | Operator’s Room Environment
20.1 | Temperature range 20-26° C (68-79° F)
20.2 | Relative air humidity without condensation (recommended level 30-70%)
20.3 | HVAC requirement removal of 3.1 kW (10,600 BTU/hr)
20.4 | Recommended room size ≥3.2 m (10.5 ft) x ≥4 m (13 ft)
21 | System Power Requirements
21.1 | Nominal voltage: 3/N~ 380-480 V (±10%) in 20 V steps
21.2 | Nominal line frequency: 50; 60 Hz (+2 Hz)
21.3 | Line impedance: 80-140 mOhm (dependent on voltage)
21.4 | Maximal power consumption: 150 kVA
22 | Software Configuration–Optional
22.1 | HD•PET
22.2 | syngo Image Fusion
22.3 | syngo Security Package
22.4 | HD•Chest
22.5 | Cardiac PET
22.6 | Cardiac CT
22.7 | Cardiac PET/CT
22.8 | Dynamic PET
22.9 | SMART Auto Cardiac Registration
22.10. | SMART Neuro AC
22.11 | NEMA Test Suite
22.12 | CT SAFIRE
22.13 | CT single-source dual-energy scan mode
23 | Patient Handling System: Continuous Bed Motion
FlowMotion
Multiparametric PET
24 | Auto QC: QualityGuard
25 | Respiratory Motion Correction
25.1 | Oncofreeze
25.2 | PET Respiratory Gating
25.3 | CT Respiratory
25.4 | Gating/Triggering
25.5 | PET + CT Respiratory Gating
25.6 | Cardiac Gating: Cardiofreeze
26 | Upgradable to latest SiPM System: Forklift Upgrade