## Sheet1
Automatic Injector (PET Infusion System)
ITEM | Technical Specifications
1 | Automatic injector (PET infusion system)
1.1 | Weight approximately 330 kg
1.2 | Dimensions approximately≤ 64 cm wide, ≤ 81 cm deep, ≤ 90 cm high.
1.3 | Maximum bulk radioactivity 30GBq (811 mCi)
1.4 | maximun radiopharmaceutical volume 30mls
1.5 | Dose range1-30mci
1.6 | dose accuracy ≤ 5% of the measured dose
1.7 | system radiation profile ≤ 0.014mSv/hr at 30.5cm from the surface of the cart at maximum activity.
1.8 | flow rate ≤ 1ml/sec
1.9 | frequency 50/60 Hz
1.10. | power shall be ≤ 300VA
1.11 | VOLTAGE 110-240VAC
1.12 | maximum radiopharmaceutical concentration100mCi/ml
1.13 | Energy range: 25keV - 3 MeV
1.14 | POWER SUPPLY
1.14.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
1.14.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
1.14.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
1.14.4 | A 3 phase in / out power supply must be offered. State details.
1.14.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
1.14.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
1.14.7 | The lifetime of UPS must be at least 15 years.
1.14.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
1.14.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
1.14.10 | The mains cable of the unit tendered for must be SABS colour coded. State details.
1.14.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
1.14.12 | Maintenance free batteries must be offered.