## Sheet1
Dual Detector Stationary Gamma Camera (SPECT CT) (CT = 16 slice)
Item | Technical Specifications | Weight
1 | General Requirements
1.1 | The system is designed for general purpose whole body and planar Nuclear Medicine Studies | 5
1.2 | The system offers SPECT CT (CT = 16 slice) and GSPECT imaging | #
1.3 | The System offers ECG gating functionality | 5
2 | Detector and Detector assembly
At least one of the two detectors shall allow for caudal/cephalic tilt.
2.1 | The minimum UFOV dimensions (min 50x38cm) | 3
2.1.1 | Please state the crystal thickness ie 3/8" | 3
2.2 | Specify the number of PMTs (min 59) and their diameter in cm | #
2.3 | The maximum Count rate is at least 220000 cps | 5
2.4 | The energy range is 40 to 580 keV | 5
2.5 | Detector parameters (NEMA Standard)
Intrinsic spatial resolution (FWHM) with 20% energy window for Tc-99m
2.5.1 | FWHM, UFOV (<4mm) | 5
2.5.2 | FWTM, UFOV (<7mm) | 5
Intrinsic energy resolution for Tc-99m
2.5.3 | FWHM, 140keV (<10%) | 5
Intrinsic linearity
2.5.4 | Differential, UFOV (<0.2mm) | 5
2.5.5 | Absolute, UFOV (<0.6mm) | 5
System spatial resolution with LEHR collimator
2.5.6 | FWHM 10cm, without scatter (<8mm) | 5
Intrinsic flood field uniformity with 20% energy window for Tc-99m
2.5.7 | Differential, UFOV (<3%) | 5
2.5.8 | Integral, UFOV (<4%) | 5
Automatic corrections
2.5.9 | Energy correction | 5
2.5.10 | Linearity correction | 5
2.5.11 | Uniformity correction | 5
2.5.12 | Center of rotation correction | 5
3 | Gantry
3.1 | The Gantry contains dual detectors. Provide brief details regarding the method of gantry construction | #
3.2 | The gantry should include a persistence monitor mounted on gantry. | #
3.3 | The gantry shall support variable angle configurability of the detectors including 76°, 90° and 180°. | #
3.4 | The automatic contouring must be available | #
3.7 | Please state the distance range between both detectors | 3
3.8 | Should be able to image patients while standing or seated on wheelchairs | 3
3.9 | Specify any limitation on any for the motions in 3.2 that depend on the type of collimator | #
3.10 | The movement controls (hand controls) must be on both sides of the gantry. | #
The emergency button to stop the movements must be on both sides of the gantry
3.11 | Manual controls are available to move the detector head away from the patient in the event of a power failure. | 5
3.12 | Gantry weight (without collimators) [kg] | 3
3.13 | Typical heat dissipation [Watt/hr] | 3
3.14 | Gantry opening (diameter) must be atleast 70cm | 3
3.15 | Upright height (min & max) for both detectors | 3
3.16 | Table
The table must be motorised
It must have movement controls on both sides
The whole body scan length shall be at least 200 cm
Minimum patient table height shall be ≤ 55cm for easy loading/ unloading of patients
3.16.1 | The minimum patient weight of 220kg | 3
3.16.2 | Please state the Pallet size [length x width x thickness) [cm] | 3
3.16.3 | Attenuation correction percentage of the table palette for energy 140keV [7%] | 3
4 | Collimators
4.1 | The following Collimators are supplied with the system: | #
4.1.1 | - LEGP | 3
4.1.2 | - LEHR | 3
4.1.3 | - Pinhole | 3
4.1.4 | - HEGP | 3
4.1.5 | - MEGP | 3
Cart (or carts) for all collimators
4.2 | Describe the process of collimator exchange | 3
5 | Operator Console
5.1 | The operator Console provide the following functions: | #
5.1.1 | - Zoom Control | 2
5.1.2 | - Scaler | 2
5.1.3 | - Manual and Automatic energy selection | 2
5.1.4 | - Image rotation control | 2
Emergency stop botton
6 | Acquisition Workstation
6.1 | An independent acquisition workstation should be offerred separately from the processing workstation | #
6.2 | Customizable automatic acquisition setup allowing best workflow efficiency | 3
6.3 | A recordable DVD or CD Drive is required | 3
6.4 | Ability to acquire minimum of 2 datasets (with different settings such as matrix size, zoom, time or count settings and gating) simultaneously is preferred | 3
6.5 | Ability to launch collimator exchange automatically from acquisition menu | 3
6.6 | Number of energy windows available | 3
6.7 | User written acquisition protocols | 3
6.8 | Persistence scope monitor | 3
6.9 | System quality control tools | 3
6.10. | Static acquisitions | 3
| - Matrices (dimensions) [pixel x pixel] | 3
| - MUGA acquisitions | 3
6.11 | Dynamic acquisitions | 5
| - Maximum number of intervals (min 3) | 5
| - Matrices (dimensions) [pixel x pixel] | 5
| - Maximum rate [matrices/sec] | 5
6.12 | Whole body scintigraphy | 3
| - Maximum whole body scan length (min 190cm) | 3
| - Automatic body contouring | 3
6.13 | SPECT acquisitions | 5
| - Scanning diameter with LEHR collimators (min to max) [cm - cm] | 5
| - Detectors geometry (available angles) [min 180º, 90º, 0º] | 5
| - Detectors rotation range around camera axis (in degrees) | 5
| - Maximum rotation speed of detectors (rpm) | 5
| - ECG gated tomography | 5
| - Automatic body contouring with all types of collimators | 5
6.14 | SPECT CT
Must have the following:
image processing and analysis
image archiving-including auto archiving
image networking-including auto transfer
Auto view of newly reconstructed images
DVD writer
32 GB or better RAM for the main system
applications for dose reduction of paediatric patients must be available
Reconstruction must be multi modality (SPECT and CT)
7 | Processing Workstation
The unit must come with atleast 2 processing workstation
7.1 | An independent acquisition workstation should be offerred separated from the processing workstation and should include | #
| - Color monitor (min 19'') | 5
| - RAM memory [min 2GB] | 5
| - Hard Disk capacity [min 500GB] | 5
| - Data archiving media | 5
7.2 | Software for automatic SPECT data reconstruction iterative method | 5
7.3 | Software packages should include all routine Nuclear Medicine procedures for: | 5
| - Quantitative analysis for cardiac SPECT studies (including cardiac gated SPECT) with normal patient databases [QBS, QPS, QGS cardiac protocols] | 5
| - Planar gated data analysis | 5
| - First pass analysis | 5
| - Cardiac shunt analysis | 5
| - Quantitative analysis of lung perfusion and ventilation | 5
| - Gallbladder ejection fraction analysis | 5
| - Gastric emptying evaluation | 5
| - Comprehensive renal analysis (deconvolution, transit times, GFR, ERPF, renogram) | 5
| - Liver functin analysis | 5
| - Semi-quantitative brain perfusion analysis | 5
| - Thyroid uptake and parathyroid analysis | 5
| - Total body knitting | 5
- SPECT viewers for brain, liver, bone
- GI bleeding
7.4 | Software packages should include: | 5
| - Manual image registration tools | 5
| - Export of images in graphic formats (JPEG, BitMap, AVI) | 5
| - Microsoft Office software | 5
7.5 | Please describe any reconstruction algorithms that improve resolution in SPECT and in Planar imaging and quote as an option. Note improved resolution to below 5mm in SPECT and Below 5.5mm in Planar imaging is preferred | 5
7.6 | SPECT CT
Must have the following:
image processing and analysis
image archiving-including auto archiving
image networking-including auto transfer
Auto view of newly reconstructed images
DVD writer
32 GB or better RAM for the main system
applications for dose reduction of paediatric patients must be available
Reconstruction must be multi modality (SPECT and CT)
7.6 | Please state if customised image processing protocols can be created | 5
Full DICOM 3.0 compatibility (send/receive, print, archiving, query/retrieve, work list) with RIS/HIS and PACS connectivity
Allow fusion with PET, MR and/or CT data from external PACS access
8 | Presentation of Results
8.1 | Color laser printer for A4 format on paper | #
8.2 | Data export/import and DICOM 3.0 format (append Dicom Conformance Statement) | #
9 | Accessories
9.1 | Head and Arm holders for Whole body imaging to be included, cardiac Spect and Brain Spect | #
9.2 | Cardiac gating device | #
9.3 | QC accessories with their mounting bracket where necessasy (as per user requirements) | #
9.3.1 | 4 bar phantom (1 in the department) | #
9.3.2 | Cobalt 57 source ( 10mCi) (1 in department) | #
9.3.3 | Jaszczak SPECT
9.3.4 | Phantom CTDI phantom
9.3.5 | CT ionization chamber with appropriate eletrometer
9.3.6 | Catphan phantom
10 | Transmission Imaging
10.1 | Spatial resolution [lp/cm] | 5
10.2 | Low contrast resolution (mm) | 5
10.3 | Transaxial field of view (cm) | 5
10.4 | Axial field of view (coverage) [cm] | 5
10.5 | Number of slices with 1 tube rotation | 5
10.6 | Slice thickness (range) [mm] | 5
10.7 | Separation between SPECT and CT planes | 5
10.8 | Plan CT of NM P-scope | 5
10.9 | Breathing technique (tidal or breathhold) | 5
11 | CT scanner
11.1 | Be integrated, spiral, fully diagnostic, multi-slice, with 16 slices;
11.2 | Provide CT data collection in support of both SPECT/CT fusion for multi-modality imaging and validated CT- AC SPECT;
11.3 | 3D CT dose modulation;
11.4 | The possibility of sub 1 mm slice thicknesses in Axial Mode;
11.5 | The minimum CT slice thickness during spiral acquisition mode shall be < 1 mm.;
11.6 | Patient port diameter/aperture of at least 70 cm;
11.7 | FOV of at least 50 cm;
11.8 | Maximum voltage of at least 130kV;
11.9 | Tube current in range 20-300 mA
11.10. | Be capable of automatically matching the CT slice thickness to the SPECT slice thickness for accurate image fusion and attenuation correction
12 | X-RAY TUBE
12.1 | Anode heat capacity must be a minimum of 3 MHU. ( State anode heat capacity, higher values preferred )
12.2 | State tube replacement pricing as well as the exchange rate.
12.3 | The tube life must be guaranteed for a minimum of 3 years.
12.4 | State the maximum anode heat dissipation rate, assuming a cold start. Higher rate preferred.
13 | DETECTOR SYSTEM:
13.1 | State the detector sampling rate
13.2 | All service phantoms to be included.
13.3 | Detector geometric efficiency must be at least 80% for all modes of operation and have 99% absorption efficiency.
B - General Issues
14 | Safety Standards of the System
14.1 | The Unit should be CE marked | #
14.2 | The Unit should be manufacture according to FDA 1020.30 regulations | #
14.3 | The Unit should be built in Compliance to IEC 601-01 | #
15 | IT Integration
15.1 | The system must have the following DICOM compatibility: | #
15.1.1 | DICOM print class | #
15.1.2 | DICOM storage class | #
15.1.3 | DICOM query/retrieve | #
15.1.4 | DICOM work list | #
15.2 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
15.2.1 | Scheduled Workflow (SWF) | 5
15.2.2 | Patient Information Reconciliation (PIR) | 5
15.2.3 | Simple image and Numeric Reports (SINR) | 5
15.2.4 | A safeguard method that protects the digital modality from potential malicious software attacks within the hospital network must be included. Describe method | 5
16 | POWER SUPPLY
16.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment. | #
16.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters). | #
16.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws. | 5
16.4 | A 3 phase in / out power supply must be offered. State details.
16.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
16.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
16.7 | The lifetime of UPS must be at least 15 years.
16.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
16.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
16.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
16.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
16.12 | Maintenance free batteries must be offered.