## Sheet1
LINEAR ACCELERATOR
Item |  | Weighting
1 | GENERAL DESCRIPTION
1.1 | Dual energy photon linear accelerator (linac) with electrons. Fitted with MLC, EPID and kV on board imaging capable of doing IMRT, VMAT, IGRT,SRS, SBRT and adapative RT.
2 | TECHNICAL SPECIFICATIONS
2.1 | Compulsory Technical specifications
2.1.1 | Photons energies of 6MV, 6MV FFF, 10MV, 10MV FFF and 15MV
2.1.2 | Must have at least five standard electron energies ranging from 4 MeV to 15 MeV
2.1.3 | MLC minimum of 120 leaves (0.5 cm at isocenter) up to maximum allowed field size.
2.1.4 | EPID high end with portal and in vivo dosimetry (state detector resolution and size)
2.1.5 | On board kV imaging (retractable)
2.1.6 | Soft/virtual/Motorized/EDW wedge.
2.1.7 | Physical/motorized wedges (15 deg to 60 deg in 15deg steps)
2.1.8 | Treatment couch must take minimum of 200 kg.
2.1.9 | Treatment couch must have carbon fibre top compatible with system in use
2.1.10 | Shadow tray (if applicable)
2.1.11 | Standard electron applicators - square from 6x6 cm to 25x25 cm (including 10x10 cm) and (Optional- circular)
2.1.12 | Lasers, sagittal and 2 x side (green or red)
2.1.13 | Optional Back pointer, Isocentric check front pointer or Back pointer laser
2.1.14 | Minimum 2 high definition CCTV, to enable full vision of patient during therapy, 1 with full movements and rotation; Intercom, last man button
2.1.15 | Collision sensors/interlogs must be available
2.2 | Functional specifications
2.2.1 | 2D, 3D , IMRT, VMAT enabled ( dynamic MLC), SRS enabled with frames and immobilizations
2.2.2 | Variable gantry speed, variable dose rate and dynamic MLC
2.2.3 | At least two Treatment planning system for one linac with 2D, 3D , IMRT, VMAT enabled ( dynamic MLC), SRS enabled licenses and atleas two Oncology workstations.( one must be added for each added Linear accelerator)
2.2.4 | Remote setup - it must be possible to set the linac up from the console without the need of an operator in the treatment room (as per manufactured recommendations, state the details)
2.2.5 | Must interface with all Record & verify, oncology management system (Very important) - OMS, state number of R& V workstations (minimum 4 for one linac and 3 for each added linac).
2.2.6 | IGRT based on kV Cone beam CT, EPID based treatment verification (2D-2D imaging with MV and kV, CBCT verification and offline review.
2.2.7 | All patient immobilization devices required to cover different cancer diagnoses, for both linear accelerator and CT scanner ( Head and Neck, Thorax, Abdomen, knees and feet rest).
2.2.8 | Head and Neck frames and all supporting accessories, breast Boards, pelvis frames, Knee rest foot rest, Masks).
2.2.9 | Include all supplied accessories for immobilization with specific details.
2.2.10 | Sample mask must be 50 for each treatment site
2.2.11 | Vacuum cushion set for different patient size ( x4 in each) with pump
2.2.12 | Compulsory onsite training and 4 offsite (international training for all speciality as agreed by all parties)
2.3 | Optional extras
Specify the cost for optional items. Please note that this will form part of the total cost for the evaluation of the tender
2.3.1 | 0.25 cm MLC for SRS treatment (Build- in or removable).
3 | Dosimetry and Patient Verification equipment (separately priced)
3.1 | The following equipment's are essential for ensuring quality, compliance and safety of delivered treatment to patients. All compulsory dosimetry equipment must be provided, installed and tested for functionality on each application.
3.2 | Barometer and thermometer with measuring probe for water temperature.
3.3 | 3 absolute dosimetry ionization chambers for photons (1), electrons (1) and Small fields (1), with electrometer and calibration certificates (refer to Dosimeter specifications)
3.4 | 1-D depth adjustable water tank with chamber holders fro TRS 398 absolute dosimetry out put measurements.
3.5 | 3D Scanning tank with all accessories required for commissioning of the Linear accelerator and treatment planning systems (training must be provided before commissioning of the all units) with at least 3 licenses and software's
3.6 | Latest technology 2D array detector complete package system for advance therapy patient pre-treatment verification in three dimension (IMRT, VMAT and SRS).
3.7 | MV, kV and cone beam CT, alignment tools, phantoms for quality assurance and image quality checks with processing software and licenses.
3.8 | Daily output check system (Diodes detector system or any accepted equivalent)
3.9 | In vivo dosimetry comprehensive solution.
3.10. | Linear array detector system for relative dosimetry (Flatness and Symmetry, Energy check), with water equivalent phantom plates (2x0.1, 1x0.2, 1x0.5, and 20x1.0 cm.), water equivalent plates with insert for photon ionization chamber and electrons ionization chamber.
3.11 | Radiation survey meter
3.12 | Radiation pocket dosimeter
3.13 | SRS and SBRT patient based pre-treatment verifications, complete solution.
3.14 | Two boxes of EBT 3 Film for film dosimetry
4 | POWER SUPPLY
4.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
4.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
4.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
4.4 | A 3 phase in / out power supply must be offered. State details.
4.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
4.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
4.7 | The lifetime of UPS must be at least 15 years.
4.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
4.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
4.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
4.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
4.12 | Maintenance free batteries must be offered.
5 | DOSIMETRY (optional extra, must be priced separatly)
5.1 | Therapy Electrometer
5.1.1 | Therapy and diagnostic dosimeter, for dose and dose rate measurements. Connecting system M type. Includes RS232, rechargeable batteries and mains cord for 115/230V, 50-60Hz
5.2 | Therapy Ionization Chambers
5.2.1 | Farmer type chamber 0.6 cm³, waterproof acrylic/aluminum; fully guarded therapy chamber with acrylic wall and aluminum electrode. Cable length 1.3m, connecting system M type. Includes acrylic build-up cap
5.2.2 | Advanced Markus electron chamber 0.02 cm³ wide guard ring, parallel plate, poly-ethylene therapy chamber with graphite coated acrylic electrode. Open chamber volume. Chamber is waterproof when using 0.87 mm acrylic cap. Cable length 1.05 m, connecting system M type
5.3 | Supporting accessories
5.3.1 | Extension cable 'M type 6m HD'
5.3.2 | 6 m 'heavy duty' detector to dosemeter extension cable, connecting system M type.
5.3.3 | Wall mount for M type connector, single
5.4 | Environmental conditions
5.4.1 | Barometer, digital precision barometer for measurement of absolute air pressure,
5.4.2 | Precision thermometer with probe -20°C to 50°C, relative humidity 10% to 95%, software Smart Graph 3, USB, LAN
5.5 | Radiation Beam Characteristics
5.5.1 | 3D water phantom, complete for atleast 50 x 50 x 41 cm scans. Includes lifting carriage/reservoir, dual channel electrometer, Two 3D chambers for measurements and reference, software with all options option. Connecting system M type. Laptop must be included
5.5.2 | 4D Pretreament verirfication system for IMRT/VMAT/SRS dosimetry. Automatic motorized tracing of the detector following the rotation of the gantry. Includes all software package and trolley. Laptop must be included
5.6 | Radiation Safety equipments
5.6.1 | Portable Survey meter, dose rate measuring with internal proportional counter tube for Gamma, 30 keV - 4.4 MeV H*(10). Range of indication 10nSv/h-100 mSv/h. Calibratable 0.5 μSv/h-100 mSv/h. Includes batteries, carrying strap, carrying bag and protective foil
5.6.2 | EBT3 Films
5.6.3 | Electronic Personal Dosemeter with alarm function and dose rate indication for the detection of the personal - dose when handling ionising radiation. :
5.6.4 | Accredited laboratory approval dosemeter for measuring gamma and X-rays for the measuring size Hp(10) (personal dose equivalent)
5.6.5 | Detector: energy compensated GM tube Radiation in front of the user's body preset dose and dose rate alarm thresholds Measuring Areas
5.6.6 | Dose indication range: 0.1 µSv ≤ Hp(10) ≤ 9.0 Sv
5.6.7 | Dose rate indication range: 0.1 µSv/h to 1.5 Sv/h.
5.7 | Radiation output measurement devices and QA
5.7.1 | 1D Scanning Arm, Tank, with Software
5.7.2 | Water Phantom - with vertical lines for alignment to isocenter using room lasers Fill lines for 20cm and 25cm depth measurements.
5.7.3 | Integral drain valve with included draining tube.
5.7.4 | Control Arm - Rotary encorder directly mounted to leadscrew shaft.
5.7.5 | 0.05mm accuracy of position and repeatability over 275.00mm travel distance.
5.7.6 | 0.05mm accuracy of position and repeatability over 275.00mm travel distance.
5.7.7 | Remote Handheld Controller.
5.7.8 | Internal memory.
5.7.9 | Software included for fully automated programmable PDD collection
5.8 | Daily Output check
5.8.1 | QUICKCHECK with docking station, for daily QA on LINACs.
5.8.2 | Measures CAX dose, dose rate, flatness, symmetry, radiation quality.
5.8.3 | Stores data in non-volatile memory for later download to PC.
5.8.4 | Features TCP/IP, USB and RS232 interfaces.
5.8.5 | Docking station and software included
5.9 | Imaging Qa
5.9.1 | EPID QC phantom QC phantom for tests on conventional EPIDs and Flat Panel Imagers.
5.9.2 | Checks linearity, isotropy, noise, low and high contrast resolution.
5.9.3 | Test elements are focused to 100 cm focus distance.
5.9.4 | For 4-25 MV photon beams.
5.9.5 | Incl. carrying case
5.9.6 | kV QC phantom QC phantom for tests on conventional Flat Panel Imagers.
5.9.7 | CathPhan Phantom