## Sheet1
3D C-Arm Flat Panel (7kW)
ITEM | TECHNICAL SPECIFICATIONS
1 | GENERATOR
1.1 | Generator must be of medium/high frequency type and have an 7kW output.
a. Radiography 110kV at 16 – 20 mA
b. Fluoroscopy a minimum of 110kV and 0.5 – 3 mA
c. State generator type:
1.2 | Automatic dose control regulating both kV and mA must be included, but manual override to be available.
1.3 | Fixed or rotating anode system may be offered. Bidder to state type.
1.4 | Anode heat capacity should be at least 40 kHU.
1.5 | X-ray tube focal spots must not be more than 0.6mm x 0.6mm for fluoroscopy and radiography.
1.6 | A High mode pulsed fluoroscopy must be possible for larger patients
1.7 | Laser positioning device mounted on generator.
1.8 | Collimation
a. System to include circular and or slot diaphragms that are able to rotate. Facility to position the diagphram without radiation must be present to reduce exposure,
b. Total filtration of x-ray tube and collimator should be at least 3mm AI.
c. Dose area product (DAP) meter must be supplied.
1.9 | System must be easily maneuverable
2 | Flat Panel Detector System
2.1 | Grid requirement: Bidder to specify type.
2.2 | Bidder to state resolution of system in Lp/cm
2.3 | Panel size of minimum 30cm x 30cm
3 | User Interface (UI)
3.1 | It must be premised on an End-User oriented design and must comply with the following:
3.2 | Touch screens with intuitive icons
3.3 | Emergency stop button
3.4 | Radiation lamp
4 | C-arm mobile stand
4.1 | Visual display of radiation warning signs.
4.2 | 3D acquisition must be motorized
4.3 | Mass of no more than 500kg
4.4 | Motor driven vertical travel more than 40cm
4.5 | Movements of the C-arm must be counter balanced and the locking system must have electromagnetic or mechanical brakes.
a. Orbital 90°both directions, Bidder to state range
b. Angulation of C-arm must be at least 190°
c. Swivel of C-arm 10° both directions
d. Movement of C-arm away from control desk at least 20cm.
Colour coding of C-arm movements.
4.6 | Stand castors must have cable deflectors
4.7 | Steering handle controlling castors must be incorporated and allow for rotation through 90°
4.8 | A central floor lock/brake for stand
4.9 | Vertical free space to be at least 70cm
4.10. | Must come with at least 2 footswitch pedals
4.11 | The whole C-arm system must be able to through a standard door entrance.
5 | Monitor cart
5.1 | Monitors
a. Dual flat screen- high definition
b. Minimum size of 19 inch touch screen
c. Connectors for external monitors
5.2 | Control cabinet of no more than 200kg
5.3 | Timer indicating lapsed screening time
5.4 | Integrated thermal printer to be supplied: Bidder to state make, model and size of prints
5.5 | 3D technology hardware and software:
a. Control software for 3D visualisation and reconstruction
b. 3D image acquisition
5.6 | Digital image processing:
a. Digital rotation of image
b. Grayscale inversion
c. Last image hold
d. Orthopaedic software (anatomic/3D/patient/metal)
e. Metal artifacts reduction software
f. Continuous and pulsed mode fluoroscopy
5.7 | Digital storage:
a. Storage capacity minimum of 4TB: Bidder to state amount
b. USB port
c. DICOM networking and software sending to PACS
5.8 | It must be compatible with navigation
6 | RADIATION PROTECTION
6.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
6.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
6.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
6.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
6.5 | A radiation report must be automatically generated and transferable via DICOM.
6.6 | Should be FDA or CE approved product
6.7 | Electrical safety conforms to standards for electrical safety IEC-60601
6.8 | The unit must comply to IEC-61267
6.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
6.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
6.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
6.12 | Lead Glass window to be installed as per room requirement
6.13 | Safety aspects of Radiation dosage leakage should be spelt out
6.14 | DB boards and electrical filters to be installed
6.15 | QA TOOLS
6.16 | Display Exposure index values
6.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
6.18 | Patient Conventional face shield to be included.
6.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
6.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
6.21 | Interlocks to stop system movement when compression is applied must be provided.
6.22 | Interlock for premature release of the X-ray switch must be available
6.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
6.24 | Phantom for calibration should be provided.
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | AIR CONDITIONER
8.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
8.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
8.3 | The air conditioner must also be included in the maintenance plan.
8.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
8.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
8.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
8.7 | Air-conditioning must be included in Maintenance plan
9 | Software Package options:
a. Orthopaedic software (anatomic/3D/patient/metal)
b. Vascular software (anatomic/3D/patient)