## Sheet1
DIGITAL FIXED CHEST X-RAY UNIT, FLOOR MOUNTED X-RAY TUBE AND ERECT BUCKY STAND WITH FIXED DETECTOR
ITEM | TECHNICAL SPECIFICATIONS
1.1 | The system must be capable of performing all general radiographic examinations of adults and children
1.2 | The system must be configured with a floor mounted tube stand on a rail and vertical stand detector combination
1.3 | The system must incorporate one flat panel digital detector that must fit into the the erect buckey :one active 43cm x 43 cm field size range .
1.4 | The system must be capable to capture a "low-energy"image and a "high-energy" image during a single examination. (Dual energy substraction chest imaging)
1.6 | System include LCD 19 inch sized, flat screen colour monitor for operators console /acquisition review station
1.7 | UPS provided to be mounted on a bracket/cabinet, not to be placed directly on the floor.
1.8 | The system must have removable grids.
1.9 | A wall mounted hanger/cabinet for the storage of removable grids and detector must be provided .
2 | X-RAY GENERATOR
2.1 | Generator should be of latest technology with high frequency and high speed starter and a power output shall be 60kW. Details are stated in supporting documents provided.
2.2 | A monitoring and display for heat status must be provided.
2.3 | AEC Functionality with 3-field ionization chamber
2.4 | APR Functionality (Anatomical programming) . Details are stated in supporting documents provided
2.5 | Maximum mA Range of 800mA and above at 100kV. Maximum mA output at 100 kV according to IEC60601-2-7 is stated.
2.6 | Selectable kV range of 40-125 kVp in 1 kVp increments
2.7 | The system must have selectable mAs. Details are stated.
2.8 | Shortest exposure time with AEC. Details are stated.
2.9 | Automatic mains compensation capability must be included.
3 | X-RAY TUBE AND FLOOR MOUNTED TUBE COLLUMN STAND
3.1 | X-ray tube should be a rotating anode high speed, compatible with digital flat panel detector. X-ray tube must be affixed to the floor mounted tube stand. Details including latest type of mode, are stated in supporting documents provided.
3.2 | Dual focal spot (small and large focus sizes) must be provided. Details of nominal focal spot sizes and focal spot combinations ,according to IEC 60336 prescription, are stated in supporting documents provided.
3.3 | Real time of tube overloading should be displayed
3.4 | Anode heat storage capacity must be 300 kHU and above. Details are stated in supporting documents provided.
3.5 | Tube overload protection is required
3.6 | Total Filtration at 70kV is 2.5mm (IEC 60601-1-3 Standard)
3.7 | Manual multileaf collimator with rotating flange collimator
3.8 | ± 45° swivelling/rotation of collimator
3.9 | Light localizer on/off switching
3.10. | Collimator must have electronic timer to terminate the light.
3.11 | X-ray Tube rotation range between ± 120 ° up to ± 135 ° . (± symbol means in both directions)
3.12 | Variable SID of the system
3.13 | Floor mounted free standing tube with rotational column design and telescopic tube arm movement, including longitudinal floor rails with 250cm travel and above.
3.14 | The longitudinal floor rail must have permanent stoppers /brackets at each end of the rail set, with shock absorbers to ensure the unit does not derail.
3.15 | Column rotation around its own axis at ± 90 ° (both directions)with a locking mechanism.
3.16 | Transverse movement
3.17 | Electromagnetic brakes with fully counter balanced mechanism
3.18 | Interlocking mechanism for 115cm /100cm and 180 cm SID . These distances must be electronically displayed on the x-ray tube /collimator housing.
4 | VERTICAL STAND WITH DETECTOR UNIT
4.1 | Vertical stand fitted with a digital flat panel - detector and three AEC measuring chambers
4.2 | Built in tray/support for 43 cm x 43 cm digital detector and removable grid
4.3 | The digital flat panel detector must be fixed . Option one
4.4 | The digital flat panel detector must be wireless and removable . Option two
4.5 | Vertical height of stand must be 2 metres and above
4.6 | Range of vertical movements of buckey: ( 1) 35cm-45 cm (minimum range height, and (2) 180cm- 200cm (maximum range height).Details are stated
4.7 | Patient hand grips for arms above head and on side of detector holder
5 | ONE X DIGITAL FLAT PANEL DETECTOR
5.1 | One x Digital FP detector of between 35cm x 43cm to 43cm x 43 cm field size.
5.2 | Amorphous silicon with cesium iodide scintillator
5.3 | Detectors have acquisition depth of 14 bits and above grayscale. Details are stated in supporting documents provided.
5.4 | Pixel pitch up to 160 micrometer (µm).
5.5 | Detector to have matrix in pixels of 2000 x 2000 and above
5.6 | For Wireless and Removable Flat Panel Detector (Option 2), include the following :
5.6.1 | Wi-Fi and LAN capability of Wireless detector . A tethered cable for wireless FP-detector option in case of Wi-Fi disruption or detector discharge .
5.6.2 | A vendor neutral wireless detector must be offered. Details are stated.
5.6.3 | The wireless detector must be able to trigger for acquisition without connectivity to the X-Ray generator to allow the wireless detector to be used in different areas.
5.6.4 | Images must be stored on the system even during battery failure.
5.6.5 | Battery pack.
5.6.6 | Battery charger.
5.6 | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
5.7 | Description of the scatter radiation grids used in conjunction with the detectors, are detailed in supporting documents provided.
5.8 | The data transmission/ reading speed from exposure to preview up to 10 seconds .
5.9 | The data transmission/ reading speed from exposure to full image up to 15 seconds .
5.10. | The flat panel battery:
5.10.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
5.10.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
5.10.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
5.10.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
5.10.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
5.10.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
5.10.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
5.10.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
5.10.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
5.11 | The Grid
5.11.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
5.11.2 | The grid must have at least:
5.11.3 | Grid ratio 8:1
5.11.4 | 100cm to 180cm focal distance
5.12 | Load capacity
5.12.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
5.12.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
6 | OPERATOR CONSOLE
6.1 | High resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor
6.2 | A diagnostic radiology LCD 19 inch imaging workstation based on latest flat panel detector technology for high image dynamics with excellent signal/noise ratio is required
6.3 | Application and examination database sotfware included.Details are stated
6.4 | Image processing software are included
6.5 | 64 bit 3.2GHz i5 Core Dual Microprocessor or similar
6.6 | RAM: 16-32GB
6.7 | 64 bit Operating system included (Windows or Similar)
6.8 | Image storage capacity up to 10 000 full format images on internal hard drive. Details of image storage capacity are stated in GB, as well as in maximum number of images in the supporting documents provided.
6.9 | No Autodeletion of images stored on internal hard drive. System to alert when near maximum storage capacity is reached.
6.10. | System must be DICOM compatible and capable of HIS/HMS/ RIS and PACS integration
6.11 | DVD /CD Writer inclusive of software for functionality of the DVD/CD Writer..
6.12 | Virus free software
F.13 | Ergonomic keyboard with mouse and function buttons
F.14 | Memory stick support/USB PORT for QC
F.15 | Touch screen monitor
7 | USER INTERFACE
7.1 | Patient data entering
7.2 | Menu
7.3 | Integrated generator operation
7.4 | Menu control
7.5 | Organ program selection
7.6 | Window position/width
7.7 | Horizontal/vertical image mirroring
7.8 | Image manipulation: rotation, cropping, contrast,brightness,sharpness, inversion
7.9 | Left/right, ap/pa working
7.10. | Configurable text annotation
7.11 | Filter selection for edge enhancement
7.12 | Filter selection for noise suppression
7.13 | Image pan and zoom capability
7.14 | measurement tools
8 | APPLICATION AND EVALUATION PROGRAMS
8.1 | Organ-related application program
8.2 | Post processing functionality
8.3 | Windowing (centre/width)
8.4 | Paging forward and back
8.5 | Multi-image display on the monitor (mosaic)
8.6 | Retro-collimation facilities (shutter)
8.7 | Image documentation in the background
8.8 | Multi-tasking technique
8.9 | Edge enhancement
8.1 | Noise suppression
8.11 | pan, zoom and image manipulation
8.12 | Individual energy substraction programs provided
9 | SYSTEM INTERPHASE
9.1 | Detector interface
9.2 | PACS ready
9.3 | Interface DICOM: "Basic Print"
9.4 | Interface DICOM: "Modality Worklist"
9.5 | Interface DICOM: "Storage"
9.6 | Interface DICOM: "Send as SCU"
9.7 | Interface DICOM "Image export/ transfer/query/retrieval to PACS/HIS/RIS
9.8 | Conformance Statement for all DICOM interfaces . Conformance Statements are included.
9.9 | Firewall protection
9.10. | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades .
12 | OPTIONAL ITEMS
12.1 | 1 X TB HARD DRIVE(external) ; DICOM 3 CONFIGURED TO STORE, TRANSFER AND PRINT X-RAY IMAGES.
12.2 | Motorized collimation control
12.3 | Tilting buckey
12.4 | Synchronized tube and detector movements. auto detector tracking and autotube positioning
12.5 | The system must have the ability to perform digital long length stitching images(scoliosis/long bones)
12.6 | Auto image stitching program included in the software package.
12.7 | TOR Rad or similar Image Resolution QC Test Tool
12.8 | A trolley with lockable castors, radiolucent top and an integrated detector holder that must move across the length of the trolley.
13 | RADIATION PROTECTION
13.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
13.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
13.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
13.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
13.5 | A radiation report must be automatically generated and transferable via DICOM.
13.6 | Should be FDA or CE approved product
13.7 | Electrical safety conforms to standards for electrical safety IEC-60601
13.8 | The unit must comply to IEC-61267
13.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
13.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
13.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
13.12 | Lead Glass window to be installed as per room requirement
13.13 | Safety aspects of Radiation dosage leakage should be spelt out
13.14 | DB boards and electrical filters to be installed
13.15 | QA TOOLS
13.16 | Display Exposure index values
13.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
13.18 | Patient Conventional face shield to be included.
13.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
13.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
13.21 | Interlocks to stop system movement when compression is applied must be provided.
13.22 | Interlock for premature release of the X-ray switch must be available
13.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
13.24 | Phantom for calibration should be provided.
14 | POWER SUPPLY
14.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
14.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
14.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
14.4 | A 3 phase in / out power supply must be offered. State details.
14.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
14.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
14.7 | The lifetime of UPS must be at least 15 years.
14.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
14.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
14.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
14.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
14.12 | Maintenance free batteries must be offered.
15 | AIR CONDITIONER
15.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
15.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
15.3 | The air conditioner must also be included in the maintenance plan.
15.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
15.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
15.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
15.7 | Air-conditioning must be included in Maintenance plan