## Sheet1
DIGITAL MAMMOGRAPHY UNIT
ITEM | GENERAL
FULL FIELD DIGITAL MAMMOGRAPHY UNIT WITH C-ARM ASSEMBLY AND ERGO DYNAMICALLY ACQUISITION WORKSTATION WITH 2D IMAGING AS WELL AS 3D DIGITAL BREAST TOMOSYNTHESIS SYSTEM
COMPUTER AIDED DIAGNOSTIC SYSTEM WITH SOFTWARE
UPRIGHT STEREO-TAXY BIOPSY SYSTEM
VACCUUM ASSISTED BREAST BIOPSY AND EXCISION SYSTEM
BIOPSY CHAIR
EXTERNAL HARD DRIVE minimum 4 TB (as option)
TECHNICAL SPECIFICATIONS
1 | FULL FIELD DIGITAL MAMMOGRAPHY UNIT WITH C-ARM ASSEMBLY AND ERGO DYNAMICALLY ACQUISITION WORKSTATION WITH 2D IMAGING AS WELL AS 3D DIGITAL BREAST TOMOSYNTHESIS SYSTEM
1.1 | GENERATOR
1.1.1 | The generator must be of Constant Potential High frequency Inverter Type
1.1.2 | The generator rating must be 7.0kW, 200mA @35kV
1.1.3 | The KV must range from 20KV to 50KV (State Range)
1.1.4 | KV increments to be 1kV
1.1.5 | The mAs must range from 3 to 400mAs (State Range)
1.1.6 | Automatic exposure device to be capable of 7 manual positions from nipple to chest wall and Auto AEC to find the optimum area of the breast
1.1.7 | The mA range must be 10 to 200mA for Large focal Spot and 10 to 50mA for small focal Spot
1.2. | X-RAY TUBE
1.2.1 | Tungsten rotating biangular anode X-ray Tube must be provided.
1.2.2 | A dual focal spot X-ray Tube must be provided at most 0.13 mm small focal and 0.33 mm large focal spot
1.2.3 | Anode heat capacity must be at least 300KHU
1.2.4 | Total inherent filtration of the X-ray tube and shield (mm) to be at least 0,63 .
1.2.5 | Filtration to be included for Tomosynthesis. State Filter:
1.2.6 | Filter selection method must be automatic and manual
1.2.7 | Tube light beam intensity must be at least 160 (LUX)
1.2.8 | Tube heat Storage must at least be 500KHU or more
1.2.9 | Fan Speed controlled in tube housing via temperature sensor must be provided.
1.3 | COLLIMATOR
1.3.1 | Collimation to be manually or automatically selected.
1.3.2 | Automatic collimation adjustment is required when compression paddles are changed
1.3.3 | X-Ray field collimation for 18x24, 24x29 ,15x15 10x10cm and 5x10cm views to be included
1.3.4 | Collimation for Tomosynthesis must be included at 18 x 29cm
1.4 | C-ARM ASSEMBLY
1.4.1 | The System must incorporate a Biopsy Ready C-arm should an upright stereotactic unit be acquired.
1.4.2 | The system must provide indented hand spaces on the side of the gantry for natural positioning of the patient hand holds
1.4.3 | The C-arm must have a Tomosynthesis Face Shield for optimum safety to the patient during Tomosynthesis and must be removable and retractable.
1.4.4 | Range of breast platform vertical movement (from floor) must be at least from 71cm to 142cm motorized.
1.4.5 | A Streamlined Tube head and SID of 70cm must be offered to provide more workspace during interventional procedures
1.4.6 | Method of angulations must be motorised from +/- +195° to -155°. State range.
1.4.7 | Angle and SID must be displayed on the Gantry and the C-arm Head as well as on the Acquisition station
1.4.8 | Compression thickness and compression force information must be displayed on both sides of the Compression device and on the Acquisition Station.
1.4.9 | Automatic release of compression after exposure must be supplied.
1.4.10 | Emergency release of compression during power failure must be available or during Exposure abortion or failure.
1.4.11 | C-arm rotation indents must be pre-programmable at any angle for up to four indents
1.5 | DIGITAL IMAGE RECEPTOR
1.5.1 | Direct capture flat panel technology must be offered.
1.5.2 | Pixel pitch must be minimum 70x70 microns
1.5.3 | Detector must be direct conversion type, eliminating the need to convert X-rays to Light. No Scintillation scanning method of imaging acquisition must be offered.
1.5.4 | Active area should be minimum 24x29 cm
1.5.5 | Resolution should be 3328x4096 pixels
1.5.6 | High contrast limiting resolution(1p/mm) must be at least 7 lines per mm
1.5.7 | State detector calibration frequency period. Calibration period to be at least 1 week valid
1.5.8 | Calibration to be performed by the user
1.5.9 | The chest gap between detector cover and detector plate to be at maximum 4mm
1.5.10 | Image acquisition time to display image on the workstation to be less than 4 seconds.
1.5.11 | Imaging Modes must be Conventional Mammography, Tomosynthesis, Combined Conventional and Tomosynthesis under the same compression in CC and MLO positions.
1.5.12 | A Grid reducing scatter radiation drastically must be offered.(State Grid type and advantages)
1.5.13 | A Grid with Auto motorized retraction for Geometric views and Tomosynthesis must be offered
1.5.14 | Linear Response over 400:1 in X-ray exposure or more must be offered.
1.5.15 | Output image bit data must be 14 bit
1.5.16 | X-ray Exposure level at which image pixels are saturated is not less than 500mR
1.5.17 | Failed Pixels on the Detector must be fixed and the Detector Certified as usable after Pixel fixes
1.5.18 | The Detector must be powered from its own individual supply and controllers offering the detector on/off times to be managed individually to prolong the detector’s lifespan by more than 100%
1.6 | GANTRY
1.6.1 | Input Line Voltage to be 200/208/220/230/240VAC Single Phase @50Hz with 10% mains compensation
1.6.2 | A 15VA online UPS to be installed with a Line ripple Filter
1.6.3 | The Gantry must be licensed for Tomosynthesis and must be ready for Tomosynthesis upon Installation. No additional upgrades will be accepted.
1.6.4 | The Tomosynthesis must be used on its own as well as with a Combo Mode, where a conventional mammogram is taken and a Tomo sweep is done under the same compression without manually removing a grid. The grid must be automatically removed and motorized.
1.6.5 | The sweep angle for Tomosynthesis must be 15 degrees to compensate for patient movement and minimizing breast compression time, optimizing image quality and throughput per patient.
1.6.6 | The total Tomo sweeps must be done in less than 4 seconds and the dosage must be similar to that of a Conventional mammogram. State Tomo Dosage @ACR phantom of 4.5cm
1.6.7 | The Tomo sweep exposure shots have to be rapid scans of 15 projections and each exposure taken per degree to prevent overexposure of the Patient.
1.6.8 | Reconstruction of the Projections has to be done in a manner to construct 1mm slices of the breast for the full Compression thickness.
1.6.9 | An upper 5mm of reconstruction slices above the breast Skin line must be available to compensate for any paddle offsets.
1.6.11 | The Tomosynthesis software must be able to Synthesize 2D images from 3D tomosynthesis Studies or able to be upgradable for this: State Price as optional Extra:
1.6.12 | State if the software will be able to be upgraded to perform localisation hookwire and contrast enhanced breast mammography in both 2D and Tomosynthesis.
1.7 | COMPRESSION
1.7.1 | Maximum Manual compression that can be applied must be 300(N)
1.7.2 | 2 pairs of foot pedals for compression adjustment required as well as vertical C-arm height adjustment.
1.7.3 | Adjustable and displayed compressor force with Pre, Full, Dual compression selectable modes must be available. Pre compression mode from 70 to 134N must be available.
Full motorized compression from 89 to 178N must be available with a manual override up to 300N
1.7.4 | Motor driven compression paddles for 18x24, 24x30 and ‘spot’ views to be included.
1.7.5 | The Compression Device must be configurable for Fixed Paddle and Fully Adjustable Self Tilting paddle system conforming to the natural contours of the breast applying equal force from Nipple to Chest Wall.
1.7.6 | Manual and motorised compression movements must be provided
1.7.7 | The following Paddles must be provided:
1.7.7.1 | · 24x29cm Paddle
1.7.7.2 | · 18x24cm Paddle
1.7.7.3 | · Small Breast Paddle
1.7.7.4 | · Frameless Spot Paddle
1.7.7.5 | · 10cm contact Paddle
1.7.7.6 | · 10 cm magnification Paddle
1.7.7.7 | · 5-10 cm spot contact paddle(State Size)
1.7.7.8 | · 5-10cm spot magnification Paddle(State Size)
1.7.7.9 | · Removable Face Shield
1.7.7.10 | · Removable, Retractable Tomosynthesis Face Shield
1.7.7.11 | · Magnification Platform
1.7.7.12 | · 15cm Perforated Localization Paddle
1.7.8 | Automatic Motorized paddle Shift must be possible for the next view optimizing workflow and time
1.7.9 | Compression tilt must be standard as well as fully adjustable and user selectable
1.7.10 | Range of movement of the compression device , relative to the breast support platform must be at least 30cm or more
1.8 | MAGNIFICATION
1.8.1 | Carbon Fibre magnification platform must be provided.
1.8.2 | Area of detector used for magnification must be 18x24 cm
1.8.3 | The size of the magnification compression paddles in cm must be 18x24cm Paddle,10cm x15cm Magnification paddle and 5-10cm spot magnification Paddles.
1.8.4 | Clearance above the magnification device when fitted in cm must be between 18 and 30 cm.
1.8.5 | Compressed breast thickness must be displayed when using magnification on the Gantry and the Acquisition Station.
1.8.6 | AEC function must be activated during magnification.
1.8.7 | Focal spot size used for magnification must be 0.1mm.
1.8.8 | System for magnification views 1.5 and 1.8 times to be included.
1.9 | RADIATION PROTECTION
1.9.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
1.9.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
1.9.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
1.9.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
1.9.5 | A radiation report must be automatically generated and transferable via DICOM.
1.9.6 | Should be FDA or CE approved product
1.9.7 | Electrical safety conforms to standards for electrical safety IEC-60601
1.9.8 | The unit must comply to IEC-61267
1.9.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
1.9.10 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
1.9.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
1.9.12 | Lead Glass window to be installed as per room requirement
1.9.13 | Safety aspects of Radiation dosage leakage should be spelt out
1.9.14 | DB boards and electrical filters to be installed
1.9.15 | QA TOOLS
1.9.16 | Display Exposure index values
1.9.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
1.9.18 | Patient Conventional face shield to be included.
1.9.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
1.9.20 | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
1.9.21 | Interlocks to stop system movement when compression is applied must be provided.
1.9.22 | Interlock for premature release of the X-ray switch must be available
1.9.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
1.9.24 | Phantom for calibration should be provided.
1.9.25 | Geometry Phantom for Tomosynthesis calibration must be provided.
1.10. | RADIOGRAPHER’S ACQUISITION STATION
1.10.1 | The Workstation must be Multi Core Intel Based CPU Technology.
1.10.2 | Minimum RAM requirements must be 32 GB
1.10.3 | Data Disk Drive must be 1TB or more
1.10.4 | A 3MP Greyscale Medical Grade LCD Display must be provided on a left or right articulating Arm
1.10.5 | A 1.2MP screen colour LCD Display must be provided as a user interface display which is built in the ergonomic station table, optimizing workflow and throughput. It is preferred that the said screen be of touch type.
1.10.6 | Quick login and access to the application software must be available. State method.
1.10.7 | A built in Automated barcode reader must be provided if required, in liaising with Worklist operations.
1.10.8 | A keyboard, Trackball and/or mouse on the workstation must be provided.
1.10.9 | Ergonomically Designed X-ray Exposure switches must be provided, eliminating repetitive motion discomfort. The system must be configurable for one or two lever activation as well as remote X-ray activation.
1.10.10 | DICOM Services must include Worklist, Print, Storage Commitment, Query/Retrieve and Modality Performed Procedure Step.
1.10.11 | Image storage of up to 9000 conventional images and 3000 Tomosynthesis images or more. State.
1.10.12 | The Acquisition Station must have its own UPS installed on the Station.
1.10.13 | All Patient details must be ordered via completed, In Progress, Pending, Scheduled and All.
1.10.14 | Advanced search facility must be available for searching the patient database.
1.10.15 | Query/Retrieve functionality must be available from PACS to the Acquisition Station.
1.10.16 | DICOM communications from the radiographer’s Acquisition workstation to the Radiologist review Station must be possible.
1.10.17 | A full Quality Control tab must be visible to perform all the required checks as per it’s due time including Automatic SNR/CNR, Artefacts, Image Quality, etc.
1.10.18 | The AEC Exposure modes must include: Auto Time, Auto kV, Auto Filter and Manual. In addition, a Tissue exposure Control mode must be available
1.10.19 | Implant software must be available to reprocess implant views for optimized image quality should the radiographer be unaware of an implant present.
1.10.20 | A full reject and repeat analysis report must be generated per radiographer
1.10.21 | The view orders have to be programmed according to each radiographer’s preferences.
1.10.22 | Output Groups must be setup to maximize the radiographer’s workflow.
1.10.23 | Printing of images must be available to DICOM printers as well as normal Printers
1.10.24 | Importing of Other vendors Images must be possible and must be IHE compliant and importation problem free. Reprocessing of the imported images must be possible.
1.10.25 | The Images must be Grey Scale DICOM Compliant
1.10.26 | The View Order must be configurable and programmable for each Radiographer’s preferences.
1.10.27 | A CDRW/DVD Writer must be available for importing or exporting of images
1.10.28 | A USB 2.0 Port must be available for importing or exporting of images and must be visible on the station outside.
1.10.29 | QA and QC Test patterns including AAPM-TG and SMPTE patterns must be available in the application.
1.10.30 | Only licence keys required to install a stereotactic device must be offered. No additional hardware and software upgrades must be necessary to upgrade to a Stereotactic Biopsy Device.
1.10.31 | All Image optimization and Radiologists preferences settings must be adjusted via software settings
Access to prio(old) and current images also exchange of annotation
1.11 | IMAGING ON ACQUISITION WORKSTATION
1.11.1 | Window width and level with fine adjustment tool must be available. State.
1.11.2 | The following image processing must be possible on the workstation:
1.11.2.1 | Image display:
1.11.2.2 | - Freely selectable screen layout
1.11.2.3 | - Windows settings (contrast and brightness setting)
1.11.2.4 | - Magnification, stepped and dynamic zoom
1.11.2.5 | - Image inversion (black/white)
1.11.2.6 | Ability to cross reference
1.11.3 | Roaming of image
1.11.4 | Image Mirrored
1.11.5 | Edge enhancement
1.11.6 | ROI
1.11.7 | AEC sensor position displayed on image
1.11.7.1 | - SNR Automatic Calculation
1.11.7.2 | - CNR Automatic Calculation as well as Base Value Programmable
1.11.7.3 | - Patient Information as DICOM Header info must be available
1.11.8 | Zoom and Crosshair
1.11.9 | Annotation:
1.11.9.1 | - Draw Oval marking
1.11.9.2 | - Measure a distance
1.11.9.3 | - Draw freehand
1.11.9.4 | - Add a text note
1.11.9.5 | - Send a Notice for the current image to the destination selected
1.11.9.6 | - Send notices for All images in the opened study to the selected destination
1.11.10 | Image evaluation:
1.11.10.1 | - Fit to Viewport
1.11.10.2 | - True Size
1.11.10.3 | - View Actual Pixels
1.11.10.4 | - Multi up Display
1.11.10.5 | - Image tiling Advance
1.11.11 | Administration:
1.11.11.1 | - The demographic patient data should be retrieved directly from a HIS/RIS system
1.11.11.2 | - The demographic patient data can be entered manually or via Worklist.
1.11.11.3 | - Barcode scanner for liaising with Patient Worklist, ensuring the correct Patient is in the Mammo Room
1.11.11.4 | - Retrieval of images from CD, DVD or PACS or USB 2.0 Port
1.11.11.5 | - The workstation should be fully DICOM compatible and include MPPS functionally
1.11.12 | Exposure Index must be visible after the exposure to assess the Index for the Image
1.11.13 | Tomosynthesis Projections must be displayed on the workstation
1.11.14 | Tomosynthesis Construction images must be displayed on the Workstation Display and must be scrollable and with Cine Mode as well.
1.11.15 | The possibility to add a view or a whole procedure must be available on the Workstation.
1.12 | RADIOLOGIST’S REPORTING STATION
1.12.1 | Two High Contrasts 21.3 Inch, 5 Mega Pixel monitors must be provided and must be approved for Mammography.
1.12.2 | Monitor size to be at least 422mm by 337mm
Monitors must have sensors installed with Built in auto calibration and QA self -diagnosis and report writing of QA/QC diagnostics.
1.12.3 | Monitor resolution to be 2048 x 2560
1.12.4 | User interface to be by mouse and keyboard.
1.12.5 | Displayed image bit depth must be 10 bits.
1.12.6 | Removable archive media must be CD/DVD and USB 2.0
1.12.7 | A PC with Data Drive, Application Drive and an Operating System Drive to be offered in RAID1 for optimal protection of Patient Data.
1.12.8 | Storage Space to be a minimum of 4 TB
1.12.9 | PC must be Windows based system with High end Dual Quad Processor @ minimum 32 GB RAM. Latest Windows, 64 Bit Operating System
1.12.10 | A UPS must be provided for the Workstation
1.12.11 | The Reviewing Software must be dedicated mammography Based Image review Software customized for radiologist use, optimizing Workflow and ease of use.
1.12.12 | DICOM services to include Print, Storage, Query /Retrieve/Auto Fetching of Priors from PACS and media export/import
1.12.13 | A Dedicated Mammo Keypad, apart from a normal keyboard must be provided for easy workflow, imaging tools and reviewing of Images via dedicated keys and specified Workflow.
1.12.14 | User preferences for scrolling mode, cine speed, default slab thickness and default initial slice must be available for Tomosynthesis reviewing
1.12.15 | Display initial hanging snapshot for a four-view Tomo study in less than 2 seconds must be possible
1.12.16 | Registration of, and instant toggle between 2D and 3D images
1.12.17 | Automatic scrolling through a limited range of slices must be possible
1.12.18 | Generate Tomosynthesis Movies as .AVI files for presentation and possible to export movies
1.12.19 | Annotate or tag reconstructed slices and store to PACS as a GSPS or DICOM Secondary Capture Image
1.12.20 | Tag reconstructed slices to save as TIFF images for presentation and printing
1.12.21 | Efficient review with personalised report flows for any clinical situation per Radiologist must be provided.
1.12.22 | Single click navigation through each programmed step must be possible
1.12.23 | Intelligent roaming through images must be possible and be programmable for each Radiologists Preferences
1.12.24 | Automatic step-through full image resolution of breast quadrants
1.12.25 | Automatic retrieval (Auto Fetching) of prior studies from PACS upon receipt of current images must be offered. State
1.12.26 | Manually or Automatically merge patient cases if patient presents with different ID and date of Birth combination must be offered. State
1.12.27 | Receive and analyze images from remote sites and mobile units through an existing PACS or SecureXchange product must be possible.
1.12.28 | Display of standard mammography images for all FFDM vendors as well as CR mammography, MRi mammography and Ultrasound MG must be possible. State
1.12.29 | Correct comparisons of current and priors images
1.12.30 | Clarified display of CAD marks and information on 2D must be possible
1.12.31 | Consistent Patient, study and image information labelling to be displayed
1.12.32 | Digital Communication channel between the radiologist and the Radiographer, independent of PACS must be possible.
1.12.33 | Exchange markings, annotations, comments, queries and instruction must be possible and must be able to save to PACS.
2 | COMPUTER AIDED DIAGNOSTICS PC AND SOFTWARE
2.1 | CAD PC with 2.6GHz Processor or more. 32 GB RAM or More, HDD space of 4 TB or more must be offered
2.2 | CD RW/DVD Drive with 2 Ethernet ports to be offered
2.3 | CAD must support up to 4 licensed digital ports
2.4 | CAD Software using sophisticated algorithms, searches digital mammograms for potential micro calcifications and masses- characteristics commonly associated with breast cancer to be offered
2.5 | The Algorithms must also incorporate anatomic correlation technology where this technology correspond findings in the contra lateral breast and different views of the breast.
2.6 | CAD must support outputs to PACS or DICOM Workstations
2.7 | Throughput of CAD cases must be 40-60 per hour/nominal
2.8 | CAD Display able to be archived to PACS
2.9 | Additional CAD specifics with Calcifications and Masses.
2.10. | CAD specificity and sensitivity to be programmed as per Radiologist’s preferences.
CAD overlay markers to be visible on Images when requested.
3 | UPRIGHT STEREO-TAXY BIOPSY SYSTEM
3.1 | BIOPSY SYSTEM
3.1.1 | Weight of the Biopsy system shall be less than 7kg
3.1.2 | Ability to do stereo Biopsy in any angle of the c-arm rotation
3.1.3 | The Biopsy window should be a minimum of 6 x 7cm.
3.1.4 | Auxiliary paddle should be included to offer a solution to biopsy in the auxiliary region
3.1.5 | Stereo pair images shall be acquired automatically without requiring the technician to move the x-ray tube.
3.1.6 | Large SID, equal to or larger then 68cm, to allow for easy installation of the biopsy system.
3.1.7 | Detachable control unit
3.1.8 | Screen color display on the control unit must be provided. It is preferred that this should be of the touch type.
3.1.9 | Biopsy add-on shall support Tomo guided biopsy procedures.
3.1.10 | Transparent Biopsy compression paddles (5cm x 5cm axillary biopsy paddle, 5cm x 5cm standard biopsy paddle and a 6cm x 7cm standard biopsy paddle).
3.2 | OPERATOR CONSOLE
3.2.1 | Biopsy user interface and targeting shall be integrated on the mammography system acquisition workstation.
3.2.2 | Target should be possible on both stereo, or on one stereo and one scout.
3.2.3 | Ability to acquire a scout with 18x24cm collimation, to be able to view tissue around biopsy window.
3.2.4 | Possibility to target up to 6 targets.
3.2.5 | Automatic placement of target on the third image after targeting on two images.
3.2.6 | Possibility to program different biopsy devices and needles.
3.2.7 | Automatic calculation of safety margins versus the detector and the paddle according to the needle used.
3.2.8 | DICOM 3.0 Print/Storage/Storage commitment and worklist approved.
3.2.9 | X-Ray generation and detector control functions integrated in a single operator console and controlled through the graphical user interface.
3.2.10 | Control device capable of tracking actual number of exposures by display icon versions of acquired images.
3.2.11 | Grayscale monitor for previewing images, 3MP minimum.
3.2.12 | Preview and image info (KV, mas, dose, compression thickness, force) Available after every exposure.
3.2.13 | Ability to add comments to an image
3.2.14 | Ability to edit views.
4 | LATERAL ARM
4.1 | 15 cm lateral biopsy compression paddle and a support paddle for lateral arm
4.2 | Should come in a case holder with needle guides and support
5 | VACUUM ASSISTED BREAST BIOPSY AND EXCISION SYSTEM MUST BE OFFERED AS AN OPTION.
5.1 | To include the following:
5.1.1 | Breast Biopsy and Excision system with
5.1.2 | Footswitch for use in MRI as well
5.1.3 | Power Cord
5.1.4 | Vacuum line and assembly with Line Filtering
5.1.5 | Stereotactic Adapter for use on system offered
5.2 | Must be compatible with the stereotactic system offered and utilized for handheld MRI.
5.3 | Must be compatible with Ultrasound use
5.4 | Using needled lengths of 9, 12 and 14 cm.
Needle gauge of 9 and 12 gauge.
5.5 | Vacuum generated: 71 cm Hg at sea level
5.6 | Biopsy and Excision system must be Mobile on rugged, Sturdy Base for use with other Modalities.
6 | MAMMOGRAPHIC BIOPSY CHAIR
6.1 | Mammography chair with low, removable arm rest and adjustable height, allowing easy access to the unit on offer, must be included.
6.2 | Must include a Foot pedal controlled hydraulic pump
6.3 | Four caster lock system to lock the wheels.
6.4 | 300lb Maximum patient weight
6.5 | Upright Seated position: 25” W x 38.5”H
6.6 | Fully reclined: 25” W x 73”L
6.7 | Overall Height: 29” raises to 36.75”
6.8 | Seat Back: 32.5” above the seat
6.9 | Wheels: 6” with locks.
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | AIR CONDITIONER
8.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
8.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
8.3 | The air conditioner must also be included in the maintenance plan.
8.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
8.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
8.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
8.7 | Air-conditioning must be included in Maintenance plan