## Sheet1
FLOUROSCOPIC SYSTEM
ITEM | SPECIFICATIONS
1 | REMOTE CONTROLLED TABLE.
1.1 | A remote controlled table with easy patient transfer access. Please state confirmation and describe table top mounting.
1.2 | The table top must be able to accommodate a patient weight of at least 200 kg or above.
1.3 | The table system must be able to tilt to +90 degree's and to -15 degree's (minimum range)
1.4 | The table top to floor height should be stated. Low lower limit is preferred.
1.5 | Table top must be X-ray translucent
1.6 | Table top length approximately 220cm x 55cm
1.7 | The table system must have a motorised compression cone device. Please state compliance.
1.8 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
1.9 | The tomographic device should have a variable layer height. State range.
1.10. | The table must be supplied with the following:
1.10.1 | Foot rest
1.10.2 | Hand grips
1.10.3 | Shoulder rests
1.10.4 | Band compression device (for IVP)
1.10.5 | Mattress with water resistant covers
1.10.6 | Compensation filter pad for peripheral studies
2 | X-RAY TUBE AND COLLIMATOR SYSTEM
2.1 | The tube and collimator must be able to vary in FFD (state range).
3 | HIGH FREQUANCY GENERATOR SYSTEM
3.1 | A 65 kWatt high frequency, microprocessor controlled x-ray generator is to be supplied together with the table system
3.2 | A matching x-ray tube unit with small and large focus distance must be provided.
3.3 | The x-ray generator must be delivered with a set of matching high tension cables.
3.4 | The x-ray generator and table system must have AEC built into the systems
3.5 | A full range of standard organ programs must be supplied. State number of programs.
3.6 | The x-ray generator must have an output ability of at least 80 kV and 800 mA.
3.7 | The x-ray generator and x-ray tube unit must have a rapid starter unit ensuring anode rotation of at least 8000 RPM.
3.8 | The system must be delivered with a dose measuring chamber to meet the new national radiation protection requirements. Please state confirmation of delivery and fitment. DAP
3.9 | The table should be of a smooth and easy to clean construction.
3.10. | An automatic collimator must be provided with full control from the table side control and the control room console.
3.11 | An automatic collimator light must be included for off radiation positioning of the patient.
3.12 | State radiation reduction techniques.
4 | CCD TELEVISION CAMERA SYSTEM WITH LARGE SIZE IMAGE INTENSIFIER AND TELEVISION MONITORS.
4.1 | A CCD television camera system must be supplied with this offer and must have a resolution of at least 1024 x 1024.
4.2 | A high resolution image intensifier of a at least an effective 38 cm input diameter must be supplied with the system.
4.3 | The image intensifier must have at least three zoom/magnification sizes. Please state sizes and compliance.
4.4 | The unit must be delivered with two flat screen TFT or LCD type monitors of at least 48 cm diagonal size. One in the control room and one in the examination room.
5 | DIGITAL IMAGING SYSTEM
5.1 | A full digital imaging system is to be offered with the system and must be totally integrated into the main system.
5.2 | Please indicated if the system can accommodate a full field flat panel detector system at a later stage. Please state cost of such upgrade in your offer
5.3 | The digital system must have digital pulsed fluoroscopy as standard. Please state the frame rates offered
5.4 | Digital radiography: Minimum of 2048 x 2048 (2 kilo Matrix), 12 bit, 0.5 to 7.5 frames/sec
5.5 | The following applications should form part of the digital imaging system (1K x 1K x 12 bits):
5.5.1 | Continuous digital fluoroscopy
5.5.2 | Pulsed fluoroscopy
5.5.3 | Digital radiography
5.5.4 | Overview of images from 4 to 16 on one monitor
5.5.5 | Digital zoom with magnification of at least 2X
5.5.6 | Windowing and contrast adjustment
5.5.7 | Image flip to the left and right as well as up and down
5.5.8 | Image rotation
5.5.9 | Digital edge enhancement
5.5.10 | Digital filtration
5.5.11 | Cine loop after a scene
5.5.12 | Electronic collimation
5.5.13 | Text annotation and free word insertion
5.5.14 | Graphic calculation of angles and relative distances
5.6 | The system storage capaciity should be at least 15 000images of 1K x 1K x 12 bits
6 | GENERAL ISSUES
SAFETY STANDARDS OF THE SYSTEM
6.1 | X-ray machine should be CE marked
6.2 | X-ray Compliance to IEC 601
7 | IT INTEGRATION
7.1 | The system must have the following DICOM compatibility:
7.1.1 | DICOM print class
7.1.2 | DICOM storage class
7.1.3 | DICOM query/retrieve
7.1.4 | DICOM worklist
7.1.5 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
7.1.6 | Scheduled Workflow (SWF)
7.1.7 | Patient Information Reconciliation (PIR)
7.1.8 | Simple image and Numeric Reports (SINR)
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
8.5 | A radiation report must be automatically generated and transferable via DICOM.
8.6 | Should be FDA or CE approved product
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan