## Sheet1
FULL BODY X-RAY SCANNER
ITEM
1 | CLINICAL PURPOSE & FUNCTION
1.1 | A real time-saving, full-body digital X-ray imaging device, it provides an excellent quality X-ray image up to 6 feet in length in just 13 seconds, emitting up to 10 times less harmful dose than regular X-ray systems.
1.2 | The full-body X-ray machine plays a significant role in the management of the trauma patient, and is an important advance in the trauma imaging.
1.3 | The effectiveness and high speed of imaging dramatically reduce the resuscitation time of patients with major injury, and allow imaging of a large number of patients in a very short time
1.4 | The detection rate for treatment, low radiation dose and speed at which the whole body can be evaluated are advantages in the primary survey of acute trauma patients.
2 | PRODUCT CONFIGURATION
2.1 | The unit supplied shall be a Slit Scanner
2.2 | The system must be able to produce a full body digital X- ray image of a patient to assist in early location of life threatening injuries and locate fractures and foreign bodies.
2.3 | It shall be possible to have the unit located near or inside an emergency or resuscitation room, for this reason the unit shall emit very low doses of x - rays scattered into the environment around the system.
2.4 | The primary environment around the unit must comply with the international Commission for Radiation Protection's recommendations for x- ray exposure to members of the general public (<1mSv of radiation dose per year per any one person). Ref : ICRP Publication No 60. ISBN 0146 -6453.
2.5 | The system shall be capable of image quality of at least 4 lp/mm for spatial resolution by imaging a standard radiological line pair tool. State the resolution modes possible and describe the type of images that are normally captured in the different resolution modes.
3 | TECHNICAL SPECIFICATION
3.1 | X -Ray Generator :
3.1.1 | Generator Rating: minimum 60 KW.
3.1.2 | The system must allow for a 2 scans ( e.g.. AP and Lateral scan view) immediately after each other without overheating. The system must have a general through rate of at least 1 patient per minute. State the duty cycle of the offered sytem.
3.1.3 | The system must be equipped with emergency stop switch.
3.1.4 | The system must have an External manual exposure release switch
3.1.5 | The system must have an Automatic Exposure Control mechanism.
3.2 | Control Console
3.2.1 | Touch screen or touch key console.
3.2.2 | Selectable exposure facility for the various anatomic studies must be available
3.2.3 | Display of selected KV and mA/mAs
3.3 | X - Ray Tube and Tube Assembly
3.3.1 | Rotating Anode x- ray tube
3.3.2 | Provide thermal cut off for avoiding tube over heating.
3.3.3 | Tube rating should be higher than the specifications of generator ratings.
3.3.4 | The tube assembly must have fully counterbalanced movements with electromagnetic locks.
3.3.5 | The longitudinal travel must be able to cover a full body scan for a standard adult patient. State longitudinal travel range.
3.4 | Diagnostic Viewing Station
3.4.1 | The X- Ray unit shall be supplied with a medical diagnostic viewing station.
3.4.2 | The unit shall allow for a full function radiological viewing manipulation and storage of an X-Ray image
3.4.3 | The viewing station shall feature 20'' ( at least) high quality LCD radiological standard monitor designed for examining X- ray images. The monitor shall :
3.4.3.1 | Have a contrast ratio of at least 700:1
3.4.3.2 | Have a brightness of at least 300Cd/m2
3.4.3.3 | Have a resolution of at least 1600x 1200 pixels
3.4.3.4 | The system shal have a local storage capacty of at least 3 months of patient's information, at 50 patient per day. The system should have a removable storage device, for example a DVD - writer for transfer of images. State the size of the internal hard drive for storage of images.
3.4.3.5 | The system must be able to produce removable CD/DVD for off - line reviewing of individual or groups of cases. Each CD/DVD shal have the viewing program loaded which is needed in order to view the images on a commercially available PC.
3.4.3.6 | A plain paper printer of suitable quality will be supplied in order to print copies of images for record keeping or review purposes.
3.5 | Patient Trolley
3.5.1 | The imaging trolley shall be suitable to transport patients directly from a helicopter or ambulance to the proposed x -ray machine.
3.5.2 | The imaging trolley shall also be suitable, to function as a patient treatment trolley in the Trauma and Emergency Unit so that a minimum of patient transfers kis necessary. It shall be equipped with : a drip stand, patient retention rails,patient retention straps, x - ray translucent mattress, equipment basket and a gas cylinder holder.
4 | IT NETWORKING AND CONNECTIVITY
4.1 | The system will be DICOM Version 3 and HL7 compliant or better
4.2 | The system must have the following DICOM compatibility:
4.2.1 | DICOM print class
4.2.2 | DICOM storage class
4.2.3 | DICOM query/retrieve
4.2.4 | DICOM work list
4.3 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
4.3.1 | Scheduled Workflow (SWF)
4.3.2 | Patient Information Reconciliation (PIR)
4.3.3 | Simple image and Numeric Reports (SINR)
4.3.4 | A safeguard method that protects the digital modality from potential malicious software attacks within the hospital network must be included. Describe method
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan