## Sheet1
Mobile C – Arm image intensifier system
ITEM | TECHNICAL SPECIFICATION
1 | X-RAY GENERATOR AND X-RAY TUBE
1.1 | The generator must be of high frequency type and have an output of at least 3kW
Radiography 40kV - 110kV at 65mA or more
Fluoroscopy 40kV - 110kV and 0mA - 15mA
State generator type
1.2 | Automatic dose control regulating both kV and mA must be included, but it must be possible to select factors manually.
1.3 | A fixed or rotating anode system may be offered. State the type of system offered.
1.4 | The anode heat capacity should be at least 40 kHU or greater.
1.5 | X-Ray Tube Focal Spots must not be more than 0.6mm x 0.6mm for Fluoroscopy and Radiography
2 | COLLIMATION
2.1 | The collimation system must include both circular and slot diaphragms that are able to rotate. Facilities for positioning the diaphragms without radiation must be present in order to reduce radiation dose to the staff
2.2 | The total filtration of the x–ray tube and collimator should be at least 3 mm Al.
State the filtration offered:
2.3 | A DAP (Dose Area Product) meter must be supplied with the system offered.
3 | IMAGE INTENSIFIER/T.V. CHAIN
3.1 | The image intensifier must be dual field with the largest field at least 31cm.
State the size of field:
State actual size coverage for imaging for each unit offered. Documentary proof of the size covered must be supplied.
3.2 | Television System must have two (2) monitors of at least 19 inches. Liquid Crystal Display (LCD) or Light Emitting Diode (LED).
3.3 | The system offered must have a Last Image Hold (LIH) facility.
Indication must be displayed on the monitor to differentiate between the live and stored image.
3.4 | The bidder must state the resolution of the system __________ Lp/cm
3.5 | The image on the monitor must be able to electronically rotate.
3.6 | A timer indicating lapsed screening time must be incorporated.
4 | MOBILE STAND
4.1 | The Mobile C-Arm and Control Cabinet must be easily maneuverable.
· C-arm mass of not more than 360 kg
· Control cabinet of not more than 200 kg
4.2 | The C-Arm should have a motor-driven height adjustment of 450mm.
4.3 | The stand must have castors with cable deflectors to facilitate movement.
4.4 | A steering handle controlling the castors must be incorporated to facilitate easy manoeuvering of the unit.
4.5 | It must be possible to rotate the castors through 90º so that the unit can be moved sideways.
A central floor lock must be provided.
4.6 | The movements of the C-Arm must be counter balanced and the locking system must have electromagnetic or mechanical brakes.
(a) Orbital 90º + 25º
(b) Angulation of C-Arm must be at least 190º
i.e. image intensifier can be used either above or below patient.
(c) Swivel of C-Arm (Wig Wag) - 10º in both directions.
(d) Movement of C-Arm away from control desk at least 200mm.
4.7 | Radius of C-Arm must be at least 600mm.
4.8 | Distance between the front of the Image Intensifier to the front of the X-Ray Collimator must be at least 700mm.
4.9 | With mobile equipment the positioning of a unit at an operating table is sometimes restricted because distance between the tube/image intensifier and the base of the unit is limited. State distance between the front of the base and a point in line with the focal spot.
Bidders must provide a diagram with markings indicating this distance.
4.10. | A grid permanently mounted on image intensifier must be provided and must be covered with perspex.
4.11 | The cables must be integrated into the “C” of the C-Arm and all surfaces must be smooth and easy to clean.
4.12 | The supplier must supply ten (10) sterile cloth covers for image intensifier and ten (10) sterile cloth for tube head.
5 | IMAGING
5.1 | Fluoroscopy must be in continuous and in pulsed mode.
5.2 | The unit must provide facilities for storage of atleast 2 TB.
5.3 | A thermal printer for producing hard copies sized 80mm x 130mm
5.4 | The printer must be integrated into the mobile cart.
5.5 | Supply details of the printer eg. make, model, type of print and size.
5.6 | A foot switch and a hand switch to control fluoroscopy must be included.
5.7 | All software for DICOM compatibility must be included as well as licensing for DICOM worklist and sending to PACS.
6 | RADIATION PROTECTION
6.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
6.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
6.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
6.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
6.5 | A radiation report must be automatically generated and transferable via DICOM.
6.6 | Should be FDA or CE approved product
6.7 | Electrical safety conforms to standards for electrical safety IEC-60601
6.8 | The unit must comply to IEC-61267
6.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
6.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
6.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
6.12 | Lead Glass window to be installed as per room requirement
6.13 | Safety aspects of Radiation dosage leakage should be spelt out
6.14 | DB boards and electrical filters to be installed
6.15 | QA TOOLS
6.16 | Display Exposure index values
6.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
6.18 | Patient Conventional face shield to be included.
6.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
6.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
6.21 | Interlocks to stop system movement when compression is applied must be provided.
6.22 | Interlock for premature release of the X-ray switch must be available
6.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
6.24 | Phantom for calibration should be provided.
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | AIR CONDITIONER
8.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
8.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
8.3 | The air conditioner must also be included in the maintenance plan.
8.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
8.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
8.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
8.7 | Air-conditioning must be included in Maintenance plan
9 | OPTIONAL EXTRAS
Orthopaedic softwares
Vascular softwares
This must be listed separately and the cost of each software package must be supplied on the schedule of optional accessories at the end of this specification.