## Sheet1
MOBILE X – RAY UNIT
ITEM | TECHNICAL SPECIFICATIONS:
The unit must be able to pass through a door of 210cm height. Spare parts must be guaranteed available for the specified life of the equipment, with a minimum of 10 years.
1 | GENERATOR:
1.1 | Compact falling load generator with an output of at least 30 kW
1.2 | The maximum mAs at various kVs must be stated
These factors will be taken into consideration in the bid adjudication. The unit must have a mA output that will make radiography with reduced exposure times possible.
State:
Maximum mAs at 125 kV
Maximum mAs at 110 kV
Maximum mAs at 100 kV
Maximum mAs at 90 kV
Maximum mAs at 80 kV
Maximum mAs at 70 kV
Maximum mAs at 60 kV
Maximum mAs at 50 kV
Maximum mAs at 40 kV
1.3 | The kV control must provide for steps of 1kV
1.4 | The bidder to state the kV range
1.5 | The cable connecting the hand switch must be at least two metres. State length
1.6 | The unit offered must incorporate automatic mains compensation
2 | X-RAY TUBE AND CABLES:
2.1 | The size of the focus must be in the range of (0.8 – 1.2)
2.2 | A pair of high voltage cables and installation cables to suit installation must be provided
2.3 | Anode heat storage capacity must not be less than 40kHU
3 | COLLIMATOR:
3.1 | A light beam diaphragm must be fitted
3.2 | The leaves must be independently adjustable in both lateral and longitudinal directions
3.3 | A rotating flange between the collimator and the tube must be supplied in order to permit diagonal collimation
3.4 | The unit is required for radiography of patients in hospital beds and a FFD of less than 100cms is frequently used. A field coverage of 43 x 43cms at a focal film distance of 80cms must be achieved.
State field coverage at 80cms. _________________
4 | TUBE STAND/MOBILE BASE:
4.1 | The unit must have a base frame with the tube parked or clipped close to the base frame for transport
4.2 | The tube stand must provide for vertical movement of the focus of the X-Ray tube up to a height of at least 1,8 metres above the floor
4.3 | The bidder to state maximum height of unit when prepared for transport
4.4 | State minimum height of tube above floor level when prepared for transport
4.5 | Angulation of the tube head must be controlled by an adequate locking device
4.6 | The unit must be suitable for radiography on a theatre table and therefore the tube when positioned to give a FFD of 100 cms must extend past the line of the base plate by an adequate distance
4.7 | If the floor to focus distance were 1800 mm (maximum) state what the maximum distance would be between the front end of the base plate of the unit and the central ray
4.8 | Bidders must supply details of the movements of the tube and tube arm/column and must state whether the tube arm can swing to the right and left (wig-wag)
4.9 | The base must be fitted with an adequate floor brake
4.1 | State the mass of the unit
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan