## Sheet1
Multislot Feeder CR Reader with Mammography
ITEM | TECHNICAL SPECIFICATION
1 | The system must be able to accept patient and imaging data.
2 | The unit must be a 19 inch touch screen system.
3 | The unit must come with a auto-stitching software
4 | The system must have a configurable auto delete function that can alert when near maximum storage capacity is reached
5 | The unit must be equipped with a barcode reader or radio chipped, which will enable the radiographer to link patient data to imaging data on a specific cassette.
6 | If an exam has been pre-ordered through the RIS (Radiology Information System), the plate reader must provide the ability to acquire appropriate patient and exam identification from the RIS via a DICOM Modality work-list.
7 | The plate reader must also provide an option for the operator to enter exam and patient identification information manually, for unscheduled exams such as emergency cases
8 | The unit must accept the following radiographic cassette/IP plates sizes: 35 X 43; 24 X 30 and 18 X 24; State the throughput of the unit for each different size of cassette/IP plate. State what other cassette/IP plate sizes are available.
9 | State the cost of the various size cassettes/IP plates on the optional schedule of accessories.
10 | The minimum throughput must be 120 (35cmX 43cm) cassettes /imaging plates or higher per hour
11 | The system must have the ability to do image control and system control, as well as patient data manipulation
12 | Imaging plates must be automatically read after insertion into the plate reader.
13 | The CR Reader unit must be connected to a dedicated CR console, providing the operator with all necessary controls in respect of window/level and image transfer and formatting to the existing PACS system and/or Dry Laser Imager via the DICOM 3 standard for reporting, distribution, printing and archiving. This console must be mounted above or alongside the CR Reader. Ability to store and print.
14 | State all radiography tooling functionality of the console offered. The system should have the capability for the provision of pre – defined text and annotation.
15 | The console described in Clause T10 above must have the ability to auto-route image and patient information to either the radiologist’s or the radiographer’s workstation as well as automatic transfer of images to a RIS/ PACS system.
16 | In the event of the Hospital network failure, it must be possible to directly print images from the CR system to a Dry Laser Imager/ Printer. The Network cost between the CR System and Dry Laser Imager/Printer must be included as part of the main bid price. The network cable must be Cat 5/6 UTP Cable.
17 | The bidder must state the image storage capacity on-line on the CR Reader System.
18 | State the time-frame required to retrieve an image from CR Reader system storage for image review.
19 | RANGE OF CR CASSETTES. (INCLUDING IMAGING PLATES)
19.1 | The following range and quantity of cassettes with imaging plates must be included in the offer. State price per each cassette/IP plate as well as a total price for the following requirement. (Moderate variation in cassette sizes will be considered)
19.1.1 | Size Quantity
(i) 24 X 30 cm 2
(ii) 18 x 24 cm 2
(iii) 35 X 43 cm 2
19.1.2 | Mammography Cassettes
Size Quantity
(i) 24 x 30cm 4
(ii) 18 x 24cm 4
19.2 | The cassettes must be bar-coded or radio chipped to facilitate the operation of the system.
19.3 | The expected plate life must be 30 000 cycles. State the expected plate life in terms of cycles.
19.4 | State what warranty is given on the number of cycles per IP plate and how this is digitally recorded. Supply details.
19.5 | State if IP plates alone can be replaced, or if the complete cassette must be replaced.
19.6 | The standard resolution must be 5 pixels /mm or better for general purpose radiography, 10 pixels /mm or better for high resolution radiography and 20 pixels /mm or better for mammogarphy.
19.7 | State the standard resolution pixels /mm for general radiography, for high resolution radiography and ultrahigh resolution for mammography.
20 | POWER SUPPLY
20.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
20.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
20.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
20.4 | A 3 phase in / out power supply must be offered. State details.
20.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
20.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
20.7 | The lifetime of UPS must be at least 15 years.
20.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
20.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
20.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
20.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
20.12 | Maintenance free batteries must be offered.
21 | AIR CONDITIONER
21.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
21.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
21.3 | The air conditioner must also be included in the maintenance plan.
21.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
21.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
21.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
21.7 | Air-conditioning must be included in Maintenance plan