## Sheet1
BLACK & WHITE ULTRASOUND UNIT (ENTRY LEVEL)
ITEM | TECHNICAL SPECIFICATION
1 | The specification calls for the supply, delivery, installation and commissioning of basic/portable type ultrasound unit, fixed to a trolley with real‑time scanning and image recording systems. State make and model offered.
2 | To be used for diagnostic investigations of patients.
3 | Depending upon the clinical requirements, trans-abdominal and/or trans-vaginal Gynaecology and Obstetric scanning, cardiac scanning, general abdominal scanning, Urological scanning, and/or Anterior fontanelle scanning may be required.
4 | The offered system shall be a real‑time black and white sector Scanner M‑Mode Cardiac Screening facilities, Thermal Printer and probe as specified.
5 | The final bid price must include the following minimum components:
5.1 | Ultrasound unit as specified. | -
5.2 | Outlets for Video Recording of Images - DICOM | -
5.2.1 | Appropriate software.
5.2.2 | 3.5 MHz convex sector probe.
5.2.3 | Probe connections X 2 | -
5.2.4 | Thermal printer. | -
5.3 | Unit fixed to mobile cart
6 | OPERATOR'S CONSOLE / TROLLEY:
6.1 | The operator's console must contain the following:
6.1.1 | Connections to the probes, thermal printer and all recordings systems.
6.1.2 | All the display monitors, keyboard and variable controls necessary for the operation of the unit.
6.1.3 | It must be possible to connect two (2) probes simultaneously.
7 | DIGITAL MEMORY AND PROCESSOR:
7.1 | The processor must provide for a minimum of 256 shades of grey imaging.
7.2 | A cine-loop memory capacity must be available
7.3 | Tissue harmonic imaging must be possible
7.4 | Supply details of the programmable features of the memory must be at least.
7.4.1 | Real time pan zoom.
7.4.2 | Magnification of frozen image.
7.5 | Image storage up to a minimum of 1000 images must be available. Measurement must be possible on stored images.
7.6 | Semi‑fixed data storage is required for patient and image identification. The number of characters available must be 40:
7.6.1 | Patient data
7.6.2 | Institution name
7.6.3 | Free text
7.7 | The unit must have a freeze‑frame facility that could be operated by a push‑button. There should be no deterioration in brightness or contrast whilst the frozen image is being viewed.
7.8 | The processor should be capable of both pre‑ and post ‑processing. A minimum of an 8 bit processor is required. Details of all such processing must be provided.
7.9 | A standard 625 line video‑output must be provided for transmitting the processed analogue signal to an image recording system.
8 | DISPLAY SYSTEMS:
8.1 | All images should be displayed for identification purposes on a sharp‑focus high-resolution non‑glare monochrome monitor. The monitor of the unit must have a diagonal size of at least 15” (inch) TFT.
8.2 | The instrument should be capable of displaying real‑time B‑Mode and M‑Mode images in grey scale.
8.3 | It must be possible to display combinations of B‑Mode and M‑Mode simultaneously.
8.4 | In addition to the real‑time image display, patient information and scan details should be provided by full alpha‑numeric keyboard, as well as automatically recording the distance between the markers.
9 | MEASURING SYSTEM:
9.1 | An electronic calliper measuring system, calibrated in centimetres range, is required with an accuracy of at least 1mm.
9.2 | The system should provide two independent calliper markers in the range of (6 - 8), superimposed on the displays and capable of being positioned at any point on the image display.
9.3 | The separation between two markers should be directly readable on the margin of the display.
9.4 | The unit must be capable of circumference, volume, ellipse and trace measurements and the answers must be displayed on the margin of the image.
9.5 | Scale markers for M‑Mode Displays should be provided 1cm apart and sub‑divisions should be stated.
9.6 | On M‑Mode there should be 1cm depth markers along the transducer axis and time markers at least every half second.
9.7 | The means of operation of the electronic calliper measuring system must be by track‑ball.
10 | SOFTWARE:
10.1 | Obstetric Software must be available:-
10.2 | To calculate foetal age and foetal weight from caliper measurements
10.3 | The software offered must be able to estimate the date of delivery from the date of the last menstrual period.
10.4 | Percentile graphs and report facilities must be obtainable from recorded data.
10.5 | Any further available obstetrical software must be stated.
10.6 | Cardiac and Vascular Studies software(In Radiology, Obstetrics, Gynaecology and Including Paediatrics)
10.7 | Software must be available for calculations pertaining to blood flow studies in:-
a) Obstetrics and Gynae: umbilical cord, uterine and ovarian pathology,
b) Abdominal and small parts.
c) Vascular: Aorta, neck, deep vein thrombosis, and superficial veins.
Software pertaining to the above-named applications must be available for obtaining the following:-
a) Display of maximum, minimum and mean mode curves with spectral Doppler.
b) Pulsatility and resistive index calculation.
c) Lumen calculations via diameter/area method.
d) Percentage stenosis based on area or diameter calculation.
e) Any other desirable features of the software package must be stated.
11 | PROBES:
11.1 | Bidders must state if exchange probes are available. If so, quote the prices offered on an exchange basis and indicate what guarantee will be offered on these probes.
11.2 | Bidders must include the one probe (3.5 MHz multi frequency curvilinear sector probe and 5MHz ) as part of the offer
11.3 | Details regarding the prices of probe and any other optional extras must be stated in the “schedule of optional extras” at the end of the specification.
12 | TRANSMITTER:
12.1 | The power output should enable tissue interfaces up to a depth of 30cm to be visualised.
13 | RECEIVER:
13.1 | The receiver should have gain controls to compensate for depths up to about 30cm.
13.2 | The bidder to state the overall system gain of the unit offered.
13.3 | Any additional features must be stated and will be considered.
14 | IMAGE RECORDING:
14.1 | Each system must be supplied with a 110 mm black and white thermal printer fixed permanently to the trolley and capable of providing good quality copies of selected frozen images displayed on the monitor screen.
15 | MOBILE CART: The unit must be attached to a sturdy trolley on castors, fitted with brakes.
16 | POWER SUPPLY:
16.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
16.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
16.3 | The UPS should not adversely affect the image quality.
16.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
17 | Optional accessories (bidder must list the probes available)
Abdominal