## Sheet1
COLOUR ULTRASOUND UNIT (ENTRY LEVEL)
ITEM | TECHNICAL SPECIFICATIONS
1 | PURPOSE:
1.1 | The system is intended to serve the hospitals for the diagnostic investigations of patients. The unit must be moderately priced and have the ability to produce high quality colour doppler images.
2 | MEDICAL USES:
2.1 | The unit must have the ability to do the following:-
2.1.1 | Produce the highest quality real-time B-Mode scans, with tissue harmonics, of all body organs and any pathology or abnormal phenomena in different shades of grey and colour.
2.1.2 | The system must have a full software package to calculate the various parameters for all ultrasound imaging as specified.
3 | SYSTEM HARDWARE:
3.1 | The system must consist of the following minimum components:
3.1.1 | Operator's console, with all controls and a full alpha-numeric keyboard.
3.1.2 | A digital memory and processor.
3.1.4 | Display monitor as specified.
3.1.5 | A high precision measuring system.
3.1.6 | A recording system as specified in line 22
3.1.7 | Pulsed wave doppler with real-time duplex mode.
4 | SYSTEM SOFTWARE:
4.1 | The unit must have software for the analysis of recorded data in respect of the following investigations:-
4.1.1 | Obstetrics.
4.1.2 | Vascular flow studies in obstetrics, and central and peripheral vascular.
4.1.3 | Paediatrics.
4.1.4 | The following software must be offered as options for future up-grading of the system:
a) Panoramic imaging
b) Free hand 3D
c) Trapezoid imaging
d) State prices of the above options under “schedule of accessories” at the end of the specification.
4.1.5 | Any other software that is presently available and not included on the offer must be stated and may be taken into consideration.
4.1.6 | The bidder must state if any other software will become available in the near future. Details of the software and cost must be listed in the schedule of optional accessories.
4.1.7 | Cardiac software and continuous wave must be quoted for and offered as an option on the schedule of optional accessories for future up-grading of the system.
4.1.8 | All version software upgrades must be supplied free of charge.
4.1.9 | The software for Dicom/PACS compatibility must be offered with worklist
5 | PROBES:
5.1 | PROBE CONNECTION:
5.1.1 | A probe holder for 3 probes and cable support must be supplied.
5.1.2 | It must be possible to connect 3 probes simultaneously
5.2 | PROBES REQUIRED:
The bid price must include the following probe. The bidder must supply details of the probe and of the cost of the probe in the schedule of accessories at the end of the specification.
5.2.1 | 3.5 MHz and 5 MHz multi-frequency convex sector probe.
6 | OPTIONAL PROBES:
6.1 | The following probes should be offered as options and the prices must be stated on the price schedule at the end of the specification.
6.1.1 | Neonatal head 5 – 7.5 MHz or higher.
6.1.2 | Linear 7.5 -10 MHz probe or higher
6.1.3 | Endo vaginal 5 – 7 MHz. or higher
6.1.4 | Cardiac Probe 5 MHz or higher
6.5 | The bidder may offer other probes as optional extras.
6.6 | Details and the cost of the probes offered must be stated on the schedule of optional accessories at the end of the specification.
6.7 | The specified frequencies of the above probes are nominal values: Variations from the central frequency of 0,5MHz each way will be acceptable.
6.8 | Mechanically driven probes, except for 3D and 4D, will not be considered.
7 | ELECTRONIC FOCUSING:
7.1 | Bidders must indicate which of the offered probes are capable of electronic dynamic focusing in which case the number of focusing zones along the depth of the field should be stated.
8 | PULSED DOPPLER:
8.1 | Transducers must be capable of switching (simultaneous) between 2-D, pulsed Doppler, continuous wave doppler and colour flow Doppler imaging in real-time (triplex mode).
9 | REPLACEMENT PROBES:
9.1 | Bidders must state if service exchange probes are available.
9.2 | If available, quote the prices offered on an exchange basis and indicate what guarantee will be offered on these probes.
10 | BIOPSY NEEDLE GUIDES:
10.1 | The bidder must indicate which of the probes quoted for can be fitted with biopsy needle guides.
10.2 | Details of such needle guides should be given as well as quotations for their supply should be stated on the schedule of optional accessories at the end of the specification.
11 | TRANSMITTER AND RECEIVER:
11.1 | The power output should enable tissue interfaces up to a depth of 30cm to be visualized.
11.2 | Variable focusing during transmission must be available.
11.3 | The range of pulse repetition rates available must be stated.
11.4 | Compensation for depth should be accomplished by slide controls on the control panel, the positions of which will indicate the Time Gain Compensation(TGC) slope, or TGC should be achieved by a control that provides for variation in near gain, slope rate and slope position.
11.5 | The frame rate will be assessed during demonstration and will be taken into consideration in the clinical evaluation of the system.
11.6 | Any additional features or offers may be stated and will be considered.
12 | OPERATOR'S CONSOLE:
12.1 | The operator's console must contain the following:-
12.2 | All the electronic equipment.
12.3 | Connections to the probes, footswitches and all recording systems.
12.4 | The facilities for an integrated writable DVD/CD drive must be available for image storage.
12.5 | The console panels must be made of rust-proof material which is resistant to marine air and is of a high quality finish. The colour must be stated.
13 | DIGITAL MEMORY AND PROCESSOR:
13.1 | The memory must use at least 8 bit technology.
13.2 | The cine loop memory capacity must be at least 1000 frames.
14 | DISPLAY:
14.1 | The instrument must be capable of displaying:-
14.2 | Real-time B-Mode scans in shades of grey - minimum 256.
14.3 | Real-time B-mode, pulsed wave, and colour Doppler images in triplex mode.
15 | MONITOR:
15.1 | The display should be on a high resolution colour monitor of at least 38cm (15”) diagonal.
15.2 | The monitor must have swivel and tilt facilities to enable optimal visualization from different angles.
16 | ANNOTATION:
16.1 | The screen must display the following information:-
16.2 | Patient's name and I.D. number.
16.3 | Date and time.
16.4 | Annotation pointer.
16.5 | All scanning parameters.
16.6 | Results of calculations.
16.7 | The number of digits available for annotation must be at least 40 or above
17 | AUDIO DOPPLER:
17.1 | The Doppler shift frequency must have an audible output that can be heard via speakers.
18 | SOFTWARE:
18.1 | Obstetric Software must be available:-
18.2 | To calculate foetal age and foetal weight from caliper measurements
18.3 | The software offered must be able to estimate the date of delivery from the date of the last menstrual period.
18.4 | Percentile graphs and report facilities must be obtainable from recorded data.
18.5 | Any further available obstetrical software must be stated.
18.6 | Cardiac and Vascular Studies software(In Radiology, Obstetrics, Gynaecology and Including Paediatrics)
18.7 | Software must be available for calculations pertaining to blood flow studies in:-
a) Obstetrics and Gynae: umbilical cord, uterine and ovarian pathology,
b) Abdominal and small parts.
c) Vascular: Aorta, neck, deep vein thrombosis, and superficial veins.
Software pertaining to the above-named applications must be available for obtaining the following:-
a) Display of maximum, minimum and mean mode curves with spectral Doppler.
b) Pulsatility and resistive index calculation.
c) Lumen calculations via diameter/area method.
d) Percentage stenosis based on area or diameter calculation.
e) Any other desirable features of the software package must be stated.
19 | IMAGE PROCESSING:
19.1 | The unit must be capable of real time and frozen pan-zoom of B-Mode and colour flow displays. Magnification factors must be at least 10 times. Any better capabilities must be stated.
20 | OPERATIONAL:
20.1 | The unit must have a freeze-frame facility that must be operated by a push-button.
20.2 | There should be no deterioration in brightness or contrast whilst the frozen image is being viewed.
20.3 | Repetitive viewing in a cine-loop with freeze frame and slow motion must be possible.
21 | MEASURING SYSTEM:
21.1 | An electronic caliper measuring system, calibrated in millimeters over the scan depth, is required with an accuracy of at least 1mm.
21.2 | The system must be capable of multiple pairs of calipers.
21.3 | The calipers must be movable to any position on the screen by the trackball. And must be in the range of (6 – 8)
21.4 | The separation between markers, area, or circumference must be displayed on the screen and be easily readable.
21.5 | Scale marker displays should be provided 1cm apart.
21.6 | Details of any additional measurement capability that is present in the system offered must be supplied.
22 | RECORDING SYSTEM:
22.1 | The system must include a black and white thermal printer with a capacity to take 110mm thrmal paper
22.2 | The printer must be integrated into the unit. The cables of the printer should be easily accessible for any changes that need to be made.
23 | POWER SUPPLY:
23.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
23.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
23.3 | The UPS should not adversely affect the image quality.
23.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
23.5 | The UPS should not adversely affect the image quality.