## Sheet1
THE ART TOP END PREMIUM DIAGNOSTIC COLOUR DOPPLER ULTRASOUND UNIT WITH PULSED WAVE DOPPLER
ITEM | TECHNICAL SPECIFICATIONS
1 | IMAGING MODALITIES
1.1 | This unit must be fully digital high end Colour Doppler Ultrasound Unit
1.2 | This unit must be latest technology, please state year in which this unit was first manufactured
1.3 | Unit should have 4 active Probe connectors with electronic switching
1.4 | The unit must offer Touch command screen for easy operation
1.5 | The unit must offer a 19” high resolution,high quality LCD colour monitor
1.6 | Panel and Touch Command Screen must be fully programmable
1.7 | Panel must be able to swivel left to right and up and down for ease of work
1.8 | The unit should offer fully Advanced Dicom software with Modality worklist included in main offer
1.9 | Unit must offer Pulse Subtraction technique
1.1 | Conventional Tissue Harmonics must be offered on all probes
1.11 | Differential Tissue Harmonics to improve spatial resolution and tissue visualization must be offered
1.12 | Multiple compounding technology to increase contrast resolution should be offered
1.13 | Tissue Specific Optimization to corrects for differences in acoustic characteristics of tissue in the body to improve lateral resolution
1.14 | B-Mode must offer single , dual and quad display
1.15 | System offered not less than 2 x 350 GB hard drive
1.17 | Unit should offer depth gain in not less than 8 steps
1.18 | Pan-zoom enlargement should be up to 12 times
1.19 | Easy programmable image processing should be in not less than 8 steps using one button to enable user to change dynamic range, persistence, contrast, enhancement, smoothing and post-processing
1.2 | Independent M-mode gain adjustment should be available.
1.21 | Frame rate of 500 frames should be offered
1.22 | Unit should not offer not less than a 202 transmitting channels
1.23 | B-Mode image on linear probe must be electronically steerable not less than 10 degrees left and right
1.24 | Unit should operate on Windows 10
1.25 | Maximum depth should not be less than 40 cm
1.26 | System dynamic range should not be less than 172dB
1.27 | Pan-zoom enlargement should be in real and frozen mode up to 2cm
1.28 | Contrast Harmonics software must be available. Please quote optionally
1.29 | The unit should offer trapezoid software capabilities on all Sector and Linear Probes
1.30. | Unit should support and be upgradeable to Panoramic software and should offer not less than 400cm in length
1.31 | Images not less than 150 000 must be able to be stored on hard drive
1.32 | Elastography software ,which displays the elasticity of tissue in real time and MicroPure software to depicts small calcifications must be included
1.33 | Automatic gain and STC control as well as velocity range and base line for Spectrum dopplerr through one touch operation must be available on unit.
1.34. | Unit must be upgradeable to display volume data obtained by another modality in synchronization with the ultrasound image. - fusion
2 | DOPPLER
2.1 | PW Doppler on all imaging probes.
2.1.1 | Max detectable velocity for PW must be stated.
2.2 | Pulse Repetition Frequency should varies between 0.5 to 24kHz
2.2.1 | Unit must come with the Continuous wave doppler
2.2.2 | Max detectable velocity must be stated.
2.3 | Independent angle steering up to ±30° for colour and Doppler must be offered on Linear Probes
2.4 | Colour pallet offering 7 types independent of B-Mode and Pulse Doppler pallet to be available.
2.5 | Sampling with Pulse Doppler should be varies between 1 – 20mm
2.6 | Display of Doppler scan modes should be offered side by side and up and down
2.7 | Unit must offer real time triplex mode
3 | COLOUR DOPPLER
3.1 | Minimum of 20 colour maps should be available for colour display.
3.2 | Colour maps must be able to be selected after frozen mode.
3.3 | Independent colour gain adjustment 1 – 16 dB in 0.5 dB steps
3.4 | Maximum colour frame rate should not be less than 307 frames /sec
3.5 | Colour flow must operate to the full depth of 40cm.
3.6 | There should be a minimum of 8 processing selections i.e. wall filters, colour averaging etc. Tendered to state the Number and list the possible selections.
3.7 | Colour maps and processing selections must be able to be
Selected by the user and stored in the presets.
3.8 | Advanced Dynamic flow should be offered for vascular recognition imaging on all transducers
3.9 | Contrast Harmonic Imaging should be available
3.10. | Directional Colour Angio should be available to display the existence of blood flow and the direction of flow movement with high sensitivity
3.11 | Unit offered must be upgradable to do high resolution 3D/ 4D applications
4 | TRANSDUCERS
4.1 | All probes required must be offered and must be included in main offer.
4.2 | All imaging transducers must be multi-frequency.
4.3 | Tissue Harmonics should be available on all probes
4.4 | The selected frequency of the imaging transducer must be displayed on the screen.
4.5 | All transducers must be digitally controlled so that continuous beam forming is possible
4.6 | It must be possible to connect 4 probes & a pencil probe simultaneously, and should be electronically selectable.
4.7 | THE FOLLOWING TRANSDUCERS MUST BE OFFERED AND INCLUDED IN THE MAIN OFFER
4.7.1 | High frequency linear probe for breast imaging from 8MHz t0 18MHz
4.7.2 | Linear probe frequency range 6MHz to 12Hz
4.7.3 | Adult cardiac probe 5 MHz
4.7.4 | Pediatric cardiac probe 7.5 MHz
4.7.5 | Convex probe 3.5 MHz
5 | MEASUREMENTS
5.1 | 16 callipers (8 channel) for measuring distance angle area volume should be offered
5.2 | Basic measurements – B-Mode : Distance , Area,Angle,Volume,Stenosis ratio
5.3 | PW/CW- mode measurements – Velocity, Acceleration, Time, Heart rate, PI , RI, S/D , Flow Volume , Doppler Trace
5.4 | Unit must offer user programmable annotation for quick labelling.
5.5 | Applications measurements should automatically be labelled on touch command screen for all applications,(B-mode and Doppler). It must automatically generate a comprehensive report.
Institutions should have the opportunity to choose their specific probe requirements.
6 | RECORDING DEVICES
6.1 | Digital Black and white thermal printer must be offered
6.2 | Unit must offer DVD/CD writer/PACS for storage/recorder for storage/recording of real-time loops and still images and USB flash drive , USB-HDD
6.3 | Recorder CD/DVD-writer should be compatible with any computer/laptop, with no additional software, images must be offered.
7 | POWER SUPPLY:
7.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
7.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
7.3 | The UPS should not adversely affect the image quality.
7.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
7.5 | The UPS should not adversely affect the image quality.