## Sheet1
MID-RANGE DIAGNOSTIC COLOUR DOPPLER ULTRASOUND UNIT WITH PULSED DOPPLER
The unit required must be a fully digital, high performance diagnostic ultrasound unit. It must be multi-purpose. It will be used for Vascular, Cardiac, Abdominal, Obstetrics & Gynaecology examinations. Unit must be a compact, mobile and hosted on a fully integrated trolley with 4 wheels.
ITEM | TECHNICAL SPECIFICATIONS
1 | IMAGING MODALITIES
1.1 | This unit must be fully digital Colour flow Doppler, Pulsed Doppler and Continouos wave Doppler
1.2 | This unit must be latest technology, please state year in which this unit was first manufactured
1.3 | Unit should have 3 active Probe connectors with electronic switching
1.5 | The unit must offer a 19 inch high resolution LCD colour monitor
1.6 | The unit must have programmable panel switches for easy operation
1.7 | The unit must be fully mobile and hosted on trolley with 4 wheels
1.8 | The unit offered should offer maximum of 40cm depth.
1.9 | Unit must have Touch Command Screen
1.1 | Tissue Harmonics must be offered on all probes for excellent image quality
1.11 | Should have a fully adjustable panel for ease of operations
1.13 | System offered should offer image memory of not less than 340 MB
1.14 | Unit should offer depth gain in not less than 8 steps
1.15 | Pan-zoom enlargement should be up to 12 times on full screen
1.16 | Easy programmable image processing should be in less than 8 steps using one button to enable user to change dynamic range, persistence, contrast, enhancement, smoothing and post-processing
1.17 | Independent M-mode gain adjustment should be available.
1.18 | Frame rate of 500 frames should be offered
1.19 | Unit should not offer not less than a 192 transmitting channels
1.2 | Post processing should be available for Gain, Dynamic range, Depth etc.
1.21 | Patient browser should offered to look at saved patients information
1.22 | Pan-zoom enlargement should be in real and frozen mode up to 2cm
1.23 | Unit should have on-board reporting system
1.24 | The unit should offer trapezoid software capabilities on all Sector and Linear Probes
1.25 | Panoramic software should be available
1.26 | Single, Dual and Quad display should be available in real time
1.27 | Please state if unit is upgradeable to 3D and 4D capabilities
1.28 | The unit should offer fully advanced Dicom/PACS software with Modality worklist included in main offer
2 | DOPPLER
2.1 | PW Doppler on all imaging probes.
2.2 | Pulse Repetition Frequency should varies between 0.5 to 24 kHz
2.2.2 | Max detectable velocity must be stated.
2.3 | Independent angle steering up to ±30° for colour and Doppler must be offered on Linear Probes
2.4 | Colour pallet offering 7 types independent of B-Mode and Pulse Doppler pallet to be available.
2.5 | Sampling with Pulse Doppler should be varies between 1 – 20mm
2.6 | Display of Doppler scan modes should be offered side by side and up and down and in quad display
2.7 | Unit must offer real time triplex mode
2.8 | Offered with continouos wave doppler
3 | COLOUR DOPPLER
3.1 | Minimum of 20 colour maps should be available for colour display.
3.2 | Colour maps must be able to be selected after frozen mode.
3.3 | Independent colour gain adjustment 1 – 16 dB in 0.5 dB steps
3.4 | Maximum colour frame rate should not be less than 387 frames p/sec
3.5 | Colour flow must operate to the full depth of 40cm.
3.6 | There should be a minimum of 8 processing selections i.e. wall filters, colour averaging etc. Tendered to state the Number and list the possible selections.
3.7 | Colour maps and processing selections must be able to be selected by the user and stored in the presets.
3.8 | Advanced Dynamic flow should be offered for vascular recognition imaging on all transducers
3.9 | Contrast Harmonic Imaging should be available
3.10. | Directional Colour Angio should be available to display the existence of blood flow and the direction of flow movement with high sensitivity
4 | TRANSDUCERS
4.1 | All probes required must be offered and must be included in main offer.
4.2 | All imaging transducers must be multi-frequency.
4.3 | Tissue Harmonics should be available on all probes
4.4 | The selected frequency of the imaging transducer must be displayed on the screen.
4.5 | All transducers must be digitally controlled so that continuous beam forming is possible.
4.6 | It must be possible to connect 3 probes & a pencil probe simultaneously, and should be electronically selectable.
4.7 | 3.75 MHz Convex probe frequency should vary between 1.9 and 6 MHz to scan large and small patients. Probe should be capable of Tissue Harmonics and Differential Tissue Harmonics to offer sharp and clear images
4.8 | Trans-Vaginal Convex Probe with frequency between 4 – 11 MHz should be offered.
4.9 | Linear probe for vascular applications, small parts with tissue harmonics from 5MHz to 11MHz
4.10. | High frequency linear probe for muscular skeletal, breast imaging from 8MHz t0 18MHz
4.11 | Multi-frequency Convex probe 6MHz with frequencies that varies between 5MHz and 9 MHz Tissue harmonics for abdominal , vascular applications for neo –natal examinations
4.12 | Pulsed Wave Doppler and Colour Doppler must be available on both probes
4.13 | Adult cardiac probe 5 MHz
4.14 | Pediatric cardiac probe 7.5 MHz
5 | MEASUREMENTS
5.1 | 16 callipers (8 channel) for measuring distance angle area volume should be offered
5.2 | 16 callipers (8 channels) hip joint should be offered
5.3 | The unit must offer storage facility for up to 2000 patient fetal growth measurement data. The information must be able to be retrievable for follow-ups.
5.4 | Unit must offer 50th percentile growth graphs of the Doppler measurements RI /PI the information saved must be retrievable for all follow-up visits.
5.5 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal Weight, HC information saved must be retrievable for all follow-up visits.
5.6 | Display of graphs must be available to compare measurements of twins, triplets in single and quad display
5.7 | Unit must offer user programmable annotation for quick labelling.
6 | RECORDING DEVICES
6.1 | Digital Black and white thermal printer must be included in the offer
6.2 | Unit must offer DVD/CD writer/PACS for storage/recorder for storage/recording of real-time loops and still images
6.3 | Recorder CD/DVD-writer should be integrated part of the unit. Software should be compatible with any computer/laptop, with no additional software,
7 | POWER SUPPLY:
7.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
7.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
7.3 | The UPS should not adversely affect the image quality.
7.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
7.5 | The UPS should not adversely affect the image quality.