## Sheet1
MOBILE ULTRASOUND UNIT
ITEM | TECHNICAL SPECIFICATIONS
1 | IMAGING MODALITIES
1.1 | Unit must offer excellent high image quality in B-Mode.
1.2 | Unit should be capable of offering a Maximum frame rate of no less than 394 frames per second
1.3 | The system must offer a Digital Continuous Beam Former to generate sharp defined ultrasound beams, to reduce artifacts and to provide more consistent overall image quality
1.4 | Unit should offer a 38-cm LCD colour monitor with a flexible arm to support LCD monitor
1.5 | Unit must have a rechargeable battery and sufficient battery life when operating the unit is operating on battery power only.
1.6 | Free programmable panel with changeable labels must be offered
1.7 | Unit should offer hard drive with a minimum memory capacity of 80GB. State maximum memory capacity.
1.8 | This unit must offer Tissue Harmonics on all probes
1.9 | The Tissue Doppler imaging software should be included
1.10. | Unit must be portable but also housed on a mobile cart.
1.11 | The system must have security functions
1.11.1 | Unit should offer entry of user name and
1.11.2 | passwords auto locking if there is no activity for security purpose of patients
1.11.3 | anti virus
1.12 | Panoramic software should be available in B-mode and colour
1.13 | Images not less than 40 000 must be able to be stored on hard drive
1.14 | Unit should allow emergency log on
1.15 | Advanced directional dynamic flow must be included
1.16 | The unit should have Programmable screensaver.
1.17 | Unit must offer handle for easy portability of equipment around hospital.
1.18 | Post processing must be available on frozen images.
1.19 | System Dynamic Range should not be less than 172 dB
1.20. | Unit should offer a minimum of 5 frequencies on all probes
1.21 | The unit must be able to offer scanning methods for Electronic Convex, Linear and phased array.
1.22 | Maximum depth should not be less than 30cm.
1.23 | Unit should offer depth gain in not less than 8 steps
1.24 | Image processing adjustments via a rotary dial with 5 combinations for optimising the best image quality for each patient’s physique (user programmable) should be available.
1.25 | Pan-zoom enlargement should be in real and frozen mode up to 2cm
1.26 | Sweep speed adjustable with a minimum of 4 steps. Independent M-mode , gain adjustment should be available
1.27 | B-mode image on linear probe must be electronically steerable.
1.28 | Grey scale level should not be less than 256 shades
1.29 | Oblique steerable B-mode must be offered on linear probes
1.30. | The unit must be connected to the SERVER, Printer, the PACS via the network and the HIS/RIS
2 | DOPPLER
2.1 | There must be Pulsed Wave Doppler on all imaging probes.
2.2 | The maximum detectable velocity for PW must be stated.
2.3 | Pulse Repetitions Frequency should varies between 1.3 to 12.5kHz
2.4 | High Pulse Repetitions Frequency should varies between 5.8 to 15.8 kHz
2.5 | State Max detectable velocity
2.6 | Independent angle steering left and right up to ±20° for colour and Doppler must be offered
2.7 | Independent gain control for PW
2.8 | Colour pallet offering 7 types independent of Pulse Doppler pallet to be available.
3 | COLOUR DOPPLER
3.1 | Maximum colour frame rate should not be less than 371 frames p/sec.
3.2 | Directional colour angio should be standard
3.3 | Tissue Doppler imaging should be standard
3.4 | Minimum of 20 colour maps should be available for colour display
3.5 | Colour maps must be able to be selected after freeze.
3.6 | Independent colour gain adjustment 1 – 16 dB in 0.5 Db steps
3.7 | Colour flow must operate to the full depth of minimum 24cm.
3.8 | There should be a minimum of 7 processing selections i.e. Wall filters, colour averaging etc. State the number and list the possible selection.
3.9 | Colour maps and processing selections must be able to be selected by the user and stored in the presets
3.10. | Quad signal processing to be available
3.11 | Parallel signal processing to be available. Bidder to provide details
3.12 | Dynamic flow on all transducers should be offered
4 | IMAGE MEMORY
4.1 | Image memory to operate in both dual and quad screen. 2D,2D/2D,2D/M, 2D/Doppler,2D/C/Doppler State format
4.2 | B-Loop facility up to 511 frames should be possible
4.3 | M/D cine review of a minimum of 128 sec should be possible.
4.4 | B-Mode image on linear transducer to be electronically steerable.
5 | TRANSDUCERS
5.1 | All probes required must be offered
5.2 | All imaging transducers must be multi-frequency
5.3 | Tissue Harmonics should be available on Convex probes
5.4 | The selected frequency of the imaging transducer must be displayed on the screen. There should be at least B-mode frequencies, THI frequencies, PW frequencies and colour frequencies.
5.5 | All transducers must be digitally controlled so continuous beam forming is possible.
5.6 | A minimum of 3 active probe plus pencil connectors must be available
5.7 | All probes must be included in the main offer
5.7.1 | Multi frequency Convex probe with tissue for abdominal , O&G ,vascular applications :3.0 MHZ to 6 MHz
5.7.2 | Pediatric curvilinear probe for premature to toddler abdomens with a 5MHz to 8MHz as well as neonatal neuro-imaging must be offered
5.7.3 | High frequency and multi frequency linear probe with tissue harmonies for small parts and muscular applications from 8MHz to 14 MHz with a trapezoid footprint.
5.7.4 | Phased array multi frequency probe with tissue harmonics for trans cranial and adult cardiac applications : central frequency 5 MHz
6 | MEASUREMENTS
6.1 | On screen measurements must be possible from real time examination.
6.2 | Basic cardiac measurements must be included
6.3 | End diastolic volume (simpson, bullet)/stroke volume, cardiac output must be possible
6.4 | Not less than 8 pairs of distance callipers must be offered for B-mode measurements. Trace, area, circumference ,ellipse and volume must be possible
6.5 | Valve orifice area measurement must be offered
6.6 | 16 callipers (8 channel) for measuring distance angle area volume should be offered
6.7 | 16 callipers (8 channels) hip joint should be offered
6.8 | The unit must offer storage facility for up to 2000 patient foetal growth measurement data. Must be able to be saved and be retrievable for follow-up visits
6.9 | The unit must offer trends on all Doppler measurements e.g. Resistive Index + Pulsitility Index of the umbilical cord, MCA, uterine arteries etc. Unit must be able to store this information for the follow-up visits
6.10. | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Foetal Weight, and HC. This information must be retrievable for all follow-up visits, to detect abnormal growth of foetus early
6.11 | Unit must offer user programmable annotation and measurement labelling to reduce scanning time.
6.12 | Display of graphs must be available to compare measurements of twins, triplets and quads
7 | RECORDING DEVICES
7.1 | Black and white thermal printer must be offered
7.2 | CD/DVD Writer/Recorder should be integrated part of the unit
7.3 | CD/DVD Recorder must be operated through panel
7.4 | High quality laser printer and desktop printer must be compatible
7.5 | Advanced DICOM/PACS with full worklist must be offered and to be included in main offer.
8 | SOFTWARE UPGRADEABILITY:
8.1 | Bidders are to state the policy with regard to future software upgrades and the costs.
8.2 | The following option should quoted separately
8.2.1 | Spatial compounding
8.2.2 | Panoramic imaging
8.2.3 | Cardiac upgrade
9 | POWER SUPPLY:
9.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
9.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
9.3 | The UPS should not adversely affect the image quality.
9.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
9.5 | The UPS should not adversely affect the image quality.