## Sheet1
Adult Advanced Cardiovascular Life Support Cardiopulmonary Resuscitation (CPR) Trainer
1 | DESCRIPTION
1.1 | Adult airway task trainer for skills training in basic to advanced airway management procedures
1.2 | Provide immediate feedback during CPR for learning while doing
2 | TECHNICAL SPECIFICATION
2.1 | Complete Adult ACLS CPR Trainer with the following:
2.1.1 | Lubricant for airway procedures
2.1.2 | IV arm skin and vein set
2.1.3 | Operating system for immediate CPR performance feedback to person doing CPR - CPR metronome (Rate, depth, rhythm)
2.1.4 | Simulator operating system including system that supports student performance monitoring by instructor and detail CPR metrics data capture
2.1.5 | Virtual patient monitor
2.1.6 | Manual defibrillation plates / pads
2.1.7 | Connectors for pacing / defibrillation cables
2.1.8 | Pacing / defibrillation cable that integrates live pacing / defibrillation
2.1.9 | Pulse oximeter probe for manikin
2.1.10 | BP cuff for manikin
2.1.11 | Manikin clothing
2.1.12 | Basic tools for assembly
2.1.13 | Portable battery pack
2.1.14 | Charge cable
2.1.15 | Drainage bag
2.2 | Must demonstrate the following Airway Features:
2.2.1 | Realistic airway anatomy to the level of the vocal cords
2.2.2 | Anatomical airway obstruction if airway not properly opened
2.2.3 | Stomach distension with aggressive bag-valve-mask ventilation
2.2.4 | IV access - to administer and drainage bag
2.3 | Must support the following Airway Procedures:
2.3.1 | Head-tilt-chin lift
2.3.2 | Jaw thrust
2.3.3 | Basic airway adjuncts: Oropharyngeal Airway (OPA) and Nasopharyngeal Airway (NPA)
2.3.4 | Supraglottic airway device insertion: Laryngeal Mask Airway (LMA)/ Laryngeal Tube Airway (LTA)
2.4 | Must support the following Respiratory Support Procedures:
2.4.1 | Bag-valve-mask ventilation with chest rise
2.5 | Must demonstrate the following Anatomical and Cardiac Features:
2.5.1 | Identification of the sternal notch and xyphoid process for determining correct hand placement for chest compressions
2.5.2 | Variable chest compliance for chest compression difficulty
2.5.3 | ECG library with appropriate peri-arrest and cardiac arrest rhythms
2.5.4 | Clinical evidence of effective CPR: ECG compression artefact and palpable carotid pulses
2.6 | Must support the following Circulatory support procedures:
2.6.1 | IV access with real-time drug and fluid administration capability
2.6.2 | Live defibrillation and synchronised cardioversion with paddles or pads (real or training) at full energy levels
2.6.3 | Live transcutaneous pacing with delivery of full energy via real or training pads with programmable threshold capture and achievement of electrical and mechanical capture
2.7 | Must provide the following CPR performance feedback:
2.7.1 | Graphic and numerical visual feedback on current evidence-based CPR performance parameters: hand position, compression-ventilation ratio, depth of compression, chest recoil, hands-off time and tidal volume delivery
2.7.2 | Detailed analysis of average CPR performance over the total time over which CPR was performed (individual and team)
2.7.3 | Immediate CPR performance graphic and numerical visual feedback for learning while doing
2.8 | Additional required features:
2.8.1 | Alteration of pupil size
2.8.2 | Patient vocal sounds associated with condition / response such as vomiting sounds, coughing and moaning
2.8.3 | Capable of communicating as the voice of patient
2.9 | Virtual Patient Monitor features:
2.9.1 | The patient monitor must integrate with the manikin operating system
2.9.2 | Touch-screen interaction with patient monitor settings and parameters
2.9.3 | Variable monitor options: AED, monitor with interactive pacing and defibrillation, and monitor only
2.9.4 | The monitor must display a selection of the following programmable parameters:
2.9.4.1 | 3-Lead ECG (Lead I, II or III)
2.9.4.2 | SpO2 – numerical and plethysmograph display
2.9.4.3 | ETCO2 with variable waveform morphology
2.9.4.4 | NIBP with manual and automatic cycling with cycle time-interval selection
2.9.4.5 | Temperature (peripheral and core)
2.1 | Software and Operating System requirements:
2.10.1 | Touch-screen monitor for interaction with monitor parameters
2.11 | System Integration requirements:
2.11.1 | Must integrate with a simulation data capturing system enabling data collection for debriefing, record keeping, reporting and research
2.11.2 | Both learner feedback system and instructor feedback with data capture system must function simultaneously
2.12 | Must come with a product user guide
2.13 | Must come in a protective case
Adult ACLS CPR Trainer - Accessories and Consumables
1 | Lubricant for airway procedures
2 | Replacement skin and veins for IV access
3 | Simulated SpO2 probe
4 | Virtual patient monitor system
5 | Simulator operating system including system that supports student performance monitoring by instructor and detail CPR metrics data capture
6 | Operating system for immediate CPR performance feedback to person doing CPR - CPR metronome (Rate, depth, rhythm)
7 | Shock absorption system
8 | Compressor for spontaneous breathing
9 | Chest skin
10 | Airway system
11 | Lung bladders
12 | Manikin clothing
13 | Battery / power pack
14 | Charge cable
15 | Drainage bag