## Sheet1
Artificial Intelligence High Fidelity Male Manikin
1 | TECHNICAL SPECIFICATION
1.1 | Complete Artificial Intelligence High Fidelity Male Manikin with the following:
1.1.1 | artificial blood
1.1.2 | lubricant
1.1.3 | artificial urine
1.1.4 | talcum powder
1.2 | The manikin must conform to the following GENERAL specifications:
1.2.1 | Must be adult life-like male size manikin
1.2.2 | Must be wireless and tetherless and must be functional for prehospital transportation scenarios, which includes Bluetooth capabilities
1.2.3 | Must be able to work off main power supply and internal power supply (batteries)
1.2.4 | All required software and hardware must be supplied for full functionality
1.2.5 | Must have case-based ultrasound capabilities for virtual or real-life monitoring (fast, e-fast, and rush protocols)
1.2.6 | Must integrate or be compatible with capture and storage of audio, visual and technical data capture for debriefing, education and research
1.2.7 | Must be available in various skin colours (light, medium, and dark tones) as per end user requirements
1.2.8 | OPTIONAL: The manikin must be supplied with all required interchangeable parts to change the manikin to a female patient if so required.
1.3 | The manikin’s skin must have lifelike features with articulation on the following areas:
1.3.1 | Neck
1.3.2 | Jaw
1.3.3 | Shoulders
1.3.4 | Elbow
1.3.5 | Wrists
1.3.6 | Hips
1.3.7 | Knees
1.3.8 | Ankles
1.4 | The manikin must also be fully articulating for placement of the following common patient positions:
1.4.1 | Supine
1.4.2 | Prone
1.5 | The manikin must conform to the following AIRWAY capabilities and specifications:
1.5.1 | Must have anatomically accurate oral cavity and airways
1.5.2 | Normal and Abnormal sounds must be audible when using spontaneous ventilations or during assisted ventilations
1.5.3 | Must be able to take advanced airways and must automatically detect advanced airway placement:
1.5.4 | Endotracheal intubation (with right mainstem intubation detection)
1.5.5 | Supraglottic airway (LMA)
1.5.6 | Oropharyngeal airway
1.5.7 | Nasopharyngeal airway
1.5.8 | Must have electronic adjustments to simulate difficult airways (e.g., Tongue edema, Laryngospasm, Pharyngeal swelling)
1.5.9 | Must include “can’t intubate can’t ventilate” scenarios
1.5.10 | Must be able to perform surgical airway techniques on the manikin
1.5.11 | Must be able to perform suctioning techniques on the above airways.
1.6 | The manikin must conform to the following BREATHING capabilities and specifications:
1.6.1 | Must have lifelike, spontaneous breathing which can be adjusted using simulation programming to be either unilateral or bilateral chest rise
1.6.2 | Must have multiple anterior and posterior auscultation sounds bilaterally
1.6.3 | Must include a library of normal and abnormal auscultation sounds
1.6.4 | Must be able to support the use of a real mechanical ventilator, with or without the requirements of an adapter with use of an external specialised lung
1.6.5 | The following minimum ventilation modes must be performable on the manikin:
1.6.5.1 | CMV
1.6.5.2 | PSV
1.6.5.3 | SIMV
1.6.5.4 | CPAP
1.6.5.5 | PEEP
1.6.6 | The manikin must have the following realistic breathing adoptions:
1.6.6.1 | Variable lung unit compliance
1.6.6.2 | Variable bilateral and unilateral bronchi resistance
1.6.6.3 | Inspiratory effort and rate
1.6.6.4 | Respiratory drive
1.6.6.5 | Real CO2 exhalation with capnography waveforms
1.6.6.6 | Auto-PEEP
1.6.7 | Manikin must make provision for chest interventions:
1.6.7.1 | Hemo/pneumothorax site
1.6.7.2 | Chest tube insertion
1.6.7.3 | Chest tube placement detection
1.6.7.4 | Bleeding, and suturing
1.6.7.5 | For needle thoracentesis:
1.6.7.5.1 | Must support needle insertion
1.6.7.5.2 | Must provide audible hiss on correct insertion
1.7 | The manikin must conform to the following CIRCULATION capabilities and specifications:
1.7.1 | Must have palpable pulses, bilaterally on the following sites:
1.7.1.1 | Carotid
1.7.1.2 | Brachial
1.7.1.3 | Radial
1.7.1.4 | Femoral
1.7.1.5 | Popliteal
1.7.1.6 | Pedal
1.7.2 | All pulses must be recorded on the manikin management program
1.7.3 | The following skin changes must be performable:
1.7.3.1 | Cyanosis
1.7.3.2 | Redness
1.7.3.3 | Pallor
1.7.4 | Must be able to use real or virtual vital signs monitoring that integrates with manikin function
1.7.5 | Must be able to support real-time intravenous fluid/infusion therapy.
1.7.6 | Must automatically recognize medication administration
1.7.7 | Must be able to perform glucose needle prick test
1.7.8 | Must be capable of performing intraosseous insertion and fluid administration
1.7.9 | Must provide a capillary refill time which is adjustable
1.7.10 | Must be capable of performing intramuscular injections
1.8 | The manikin must conform to the following NEUROLOGICAL capabilities and specifications:
1.8.1 | Adjustable blinking eyes, Pupil reaction, and eye movement
1.8.2 | Pupils must realistically adjust to light, but must be programmable
1.8.3 | Eye movement must be adjustable to neurological deficits
1.8.4 | Eyes must be able to move and follow an object/person
1.8.5 | The manikin must have the ability, through AI, to listen and realistically respond to the practitioner, according to a set scenario. Speech recognition must have the ability to adapt and learn over time, while speech must be in English.
1.8.6 | The simulation controller must have pre-recorded responses
1.8.7 | Must have the ability of moving its mouth during speech
1.8.8 | Must simulate the following:
1.8.8.1 | Closed mouth
1.8.8.2 | Open mouth
1.8.8.3 | Smiling
1.8.8.4 | Unilateral droop
1.8.8.5 | Lockjaw
1.8.9 | The neck of the manikin must be able to simulate the following:
1.8.9.1 | Rotation
1.8.9.2 | Flexion
1.8.9.3 | Extension
1.8.9.4 | Reduced movement during neck injuries
1.8.10 | The following facial expressions must be included:
1.8.10.1 | Facial droop (both left and right)
1.8.10.2 | Pain experience
1.8.10.3 | Stunned face (quizzical)
1.8.10.4 | Scared
1.8.10.5 | Smiling
1.8.10.6 | Worried (programmable)
1.8.10.7 | Anxious (programmable)
1.8.10.8 | Lethargic (programmable)
1.8.11 | One or both arms of the manikin must be able to perform the following:
1.8.11.1 | Motor reflex
1.8.11.2 | Shake hand
1.8.11.3 | Squeeze hand
1.8.11.4 | Raise arm
1.8.11.5 | Withdrawal response
1.8.11.6 | Abnormal posturing (head injury – Decorticate/Decerebrate)
1.8.12 | Must indicate realistic signs of a stroke patient
1.8.13 | Must realistically react to painful stimuli
1.8.14 | Must have appropriate movement representing seizure activity
1.8.15 | Must be able to simulate sweating and tears.
1.9 | The manikin must conform to the following GENITOURINARY capabilities and specifications:
1.9.1 | Must be able to insert urinary catheter
1.9.2 | Urinary flow must be programmable
1.9.3 | Urinary incontinence must be simulated
1.9.4 | Must be supplied with an internal reservoir for urine.
1.1 | The manikin must conform to the following GASTROINTESTINAL capabilities and specifications:
1.10.1 | Bowel auscultation field (minimum 4 fields) with a library of various sounds
1.10.2 | Visible gastric inflation
1.11 | The manikin must conform to the following TRAUMA capabilities and specifications:
2.11.1 | Must have dedicated ports for bleeding simulation
2.11.2 | Must include traumatic wound accessories (arms and legs) for the purpose of trauma management protocols
2.11.3 | Bleeding must be capable of matching blood pressure
2.11.4 | Must have an internal reservoir to store blood
2.12 | The manikin must conform to the following CARDIAC capabilities and specifications:
1.12.1 | Auscultation of heart sounds must be realistic on Aorta, Pulmonary, Tricuspid, and mitral valves of the heart
1.12.2 | Must be able to realistically place both 4 lead and 12 lead ECG wires
1.12.3 | Must be able to program and simulate cardiac injuries which include:
1.12.3.1 | Ischemia
1.12.3.2 | Infarction
1.12.3.3 | Necrosis
1.12.4 | Must include a 12-lead library or must be able to generate a 12 lead ECG through software (must be supplied)
1.12.5 | The ECG rhythm must indicate respiratory patterns
1.12.6 | Must support the following cardiac interventions:
1.12.6.1 | Defibrillation and double sequential defibrillation
1.12.6.2 | Cardioversion
1.12.6.3 | Pacing
1.12.7 | Must allow defibrillation with paddle or pad placement as either:
1.12.7.1 | Anterior / Lateral
1.12.7.2 | Anterior / Posterior
1.12.8 | Manikin and software must have a feedback mechanism on compressions (rate, depth interruption time), and respirations (excessive ventilation, rate)
1.13 | Must come with a product user guide
1.14 | Must come in a protective case
Artificial Intelligence High Fidelity Male Manikin - Accessories and Consumables
1 | artificial blood
2 | lubricant
3 | artificial urine
4 | talcum powder