## Sheet1
High Fidelity Newborn Manikin
1 | TECHNICAL SPECIFICATION
1.1 | The manikin must conform to the following GENERAL specifications:
1.1.1 | Must be a 40-week-old baby manikin
1.1.2 | Must be wireless and tetherless and must be functional for prehospital transportation scenarios, which includes Bluetooth capabilities
1.1.3 | Must be able to work off main power supply and internal power supply (batteries)
1.1.4 | All required software and hardware must be supplied for full functionality
1.1.5 | Must integrate or be compatible with capture and storage of audio, visual and technical data capture for debriefing, education and research
1.1.6 | Must be available in various skin colours (light, medium, and dark tones) as per end user requirements
1.1.7 | OPTIONAL: The manikin must be supplied with all required interchangeable parts to change the manikin to a male/female patient if so required.
1.2 | The manikin’s skin must have lifelike features with articulation on the following areas:
1.2.1 | Neck
1.2.2 | Jaw
1.2.3 | Shoulders
1.2.4 | Elbow
1.2.5 | Hips
1.2.6 | Knees
1.3 | The manikin must conform to the following AIRWAY capabilities and specifications:
1.3.1 | Must have anatomically accurate oral cavity and airways
1.3.2 | Normal and Abnormal sounds must be audible when using spontaneous ventilations or during assisted ventilations
1.3.3 | Must have normal and abnormal airway sounds that are synchronized with breathing pattern and rate
1.3.4 | Must be able to take advanced airways and must automatically detect advanced airway placement:
1.3.4.1 | Endotracheal intubation (with right mainstem intubation detection)
1.3.4.2 | Supraglottic airway (LMA)
1.3.4.3 | Oropharyngeal airway
1.3.4.4 | Nasopharyngeal airway
1.3.5 | Must have electronic adjustments to simulate difficult airways (e.g., Tongue edema, Laryngospasm, Pharyngeal swelling)
1.3.6 | Must include “can’t intubate can’t ventilate” scenarios
1.3.7 | Must be able to perform suctioning techniques on the above airways.
1.3.8 | Must be able to perform the following basic airway techniques:
1.3.8.1 | Head tilt chin lift
1.4 | The manikin must conform to the following BREATHING capabilities and specifications:
1.4.1 | Must have lifelike, spontaneous breathing which can be adjusted using simulation programming to be either unilateral or bilateral chest rise
1.4.2 | Must be able to program grunting and crying sounds
1.4.3 | Must have multiple anterior and posterior auscultation sounds bilaterally
1.4.3.1 | Must include a library of normal and abnormal auscultation sounds
1.4.3.2 | Must support the use of bag valve mask ventilations
1.4.3.3 | Must be able to support the use of a real mechanical ventilator, with or without the requirements of an adapter with use of an external specialised lung
1.5 | The manikin must conform to the following CIRCULATION capabilities and specifications:
1.5.1 | Must have palpable pulses, bilaterally on the following sites:
1.5.1.1 | Fontanelle
1.5.1.2 | Umbilical
1.5.1.3 | Brachial
1.5.2 | All pulses must be recorded on the manikin management program
1.5.3 | The following skin changes must be performable:
1.5.3.1 | Cyanosis
1.5.4 | Must be able to use real or virtual vital signs monitoring that integrates with manikin function
1.5.5 | Must be able to support real-time intravenous fluid/infusion therapy.
1.5.6 | Must be able to realistically perform phlebotomy
1.5.7 | Must be capable of performing intraosseous insertion and fluid administration
1.5.8 | Must be capable of performing intramuscular and subcutaneous injections
1.6 | The manikin must conform to the following NEUROLOGICAL capabilities and specifications:
1.6.1 | The neck of the manikin must be able to simulate the following:
1.6.1.1 | Rotation
1.6.1.2 | Flexion
1.6.1.3 | Extension
1.6.2 | Must have appropriate movement representing seizure activity
1.6.3 | Must realistically move upper limbs
1.7 | The manikin must conform to the following GENITOURINARY capabilities and specifications:
1.7.1 | Must be able to insert urinary catheter (male and female genitalia must be provided)
1.7.2 | Must be supplied with an internal reservoir for urine.
1.8 | The manikin must conform to the following GASTROINTESTINAL capabilities and specifications:
1.8.1 | Bowel auscultation field
1.9 | The manikin must conform to the following ADDITIONAL capabilities and specifications:
1.9.1 | Must be articulating and have movement
1.9.2 | Must have various options to control muscle movement or tone
1.9.3 | Color of cyanosis/hypoxia must be linked to various hypoxic events
1.9.4 | Must be able to realistically perform umbilical catheterization
1.9.5 | Must be able to place a temperature probe
1.9.6 | Must be able to realistically insert feeding tubes
1.1 | The manikin must conform to the following CARDIAC capabilities and specifications:
1.10.1 | Auscultation of heart sounds
1.10.2 | Must be able to realistically place 4 lead ECG wires
1.10.3 | Must include a 12-lead library or must be able to generate a 12 lead ECG through software (must be supplied)
1.10.4 | Manikin and software must have a feedback mechanism on compressions (rate, depth interruption time), and respirations (excessive ventilation, rate)
1.11 | Must come with a product user guide
1.12 | Must come in a protective case
High Fidelity Newborn - Accessories and Consumables
1 | Packs of 10 replacement IO bones
2 | replacement skins
3 | replacement arms
4 | replacement legs