## Sheet1
High Fidelity Paediatric Manikin
1 | TECHNICAL SPECIFICATION
1.1 | The manikin must conform to the following GENERAL specifications:
1.1.1 | Must be life-like paediatric size manikin
1.1.2 | Must be wireless and tetherless and must be functional for prehospital transportation scenarios, which includes Bluetooth capabilities
1.1.3 | Must be able to work off main power supply and internal power supply (batteries)
1.1.4 | All required software and hardware must be supplied for full functionality
1.1.5 | Must integrate or be compatible with capture and storage of audio, visual and technical data capture for debriefing, education and research
1.1.6 | Must be available in various skin colours (light, medium, and dark tones) as per end user requirements
1.1.7 | OPTIONAL: The manikin must be supplied with all required interchangeable parts to change the manikin to a male/female patient if so required.
1.2 | The manikin’s skin must have lifelike features with articulation on the following areas:
1.2.1 | Neck
1.2.2 | Jaw
1.2.3 | Shoulders
1.2.4 | Elbow
1.2.5 | Hips
1.2.6 | Knees
1.2.7 | Ankles
1.3 | The manikin must also be fully articulating for placement of the following common patient positions:
1.3.1 | Supine
1.3.2 | Prone
1.3.3 | Fowler’s position
1.4 | The manikin must conform to the following AIRWAY capabilities and specifications:
1.4.1 | Must have anatomically accurate oral cavity and airways
1.4.2 | Normal and Abnormal sounds must be audible when using spontaneous ventilations or during assisted ventilations (additional hardware to be supplied)
1.4.3 | Must be able to take advanced airways and must automatically detect advanced airway placement:
1.4.3.1 | Endotracheal intubation (with right mainstem intubation detection)
1.4.3.2 | Supraglottic airway (LMA)
1.4.3.3 | Oropharyngeal airway
1.4.3.4 | Nasopharyngeal airway
1.4.4 | Must have electronic adjustments to simulate difficult airways (e.g., Tongue edema)
1.4.5 | Must include “can’t intubate can’t ventilate” scenarios
1.4.6 | Must be able to perform suctioning techniques on the above airways.
1.4.7 | Must be able to perform the following basic airway techniques:
1.4.7.1 | Head tilt chin lift
1.4.7.2 | Jaw thrust
1.5 | The manikin must conform to the following BREATHING capabilities and specifications:
1.5.1 | Must have lifelike, spontaneous breathing which can be adjusted using simulation programming to be either unilateral or bilateral chest rise
1.5.2 | Must have multiple anterior and posterior auscultation sounds bilaterally
1.5.3 | Must include a library of normal and abnormal auscultation sounds
1.5.4 | Must be able to support the use of a real mechanical ventilator, with or without the requirements of an adapter with use of an external specialised lung
1.5.5 | Manikin must make provision for chest interventions:
1.5.5.1 | Hemo/pneumothorax site
1.5.5.2 | Chest tube insertion
1.5.5.3 | For needle thoracentesis:
1.5.5.3.1 | Must support needle insertion
1.5.5.3.2 | Must detect needle placement
1.5.5.3.3 | Must provide audible hiss on correct insertion
1.6 | The manikin must conform to the following CIRCULATION capabilities and specifications:
1.6.1 | Must have palpable pulses, bilaterally on the following sites:
1.6.1.1 | Carotid
1.6.1.2 | Radial
1.6.1.3 | Femoral
1.6.2 | All pulses must be recorded on the manikin management program
1.6.3 | Must be able to support real-time intravenous fluid/infusion therapy.
1.6.4 | Must be able to realistically perform phlebotomy
1.6.5 | Must be capable of performing intraosseous insertion and fluid administration
1.7 | The manikin must conform to the following NEUROLOGICAL capabilities and specifications:
1.7.1 | Adjustable blinking eyes, Pupil reaction, and eye movement
1.7.2 | Pupil/s must realistically adjust to light, but must be programmable
1.7.3 | The simulation controller must have pre-recorded responses
1.7.4 | The neck of the manikin must be able to simulate the following:
1.7.4.1 | Rotation
1.7.4.2 | Flexion
1.7.4.3 | Extension
1.7.5 | Must have appropriate movement representing seizure activity
1.8 | The manikin must conform to the following GASTROINTESTINAL capabilities and specifications:
1.8.1 | Bowel auscultation field
1.8.2 | Visible gastric inflation
1.9 | The manikin must conform to the following CARDIAC capabilities and specifications:
1.9.1 | Auscultation of heart sounds
1.9.2 | Must be able to realistically place 4 lead ECG wires
1.9.3 | Must be able to program and simulate cardiac injuries which include:
1.9.3.1 | Ischemia
1.9.3.2 | Infarction
1.9.3.3 | Necrosis
1.9.4 | Must include a 12-lead library or must be able to generate a 12 lead ECG through software (must be supplied)
1.9.5 | Must support the following cardiac interventions:
1.9.5.1 | Defibrillation and double sequential defibrillation
1.9.5.2 | Cardioversion
1.9.5.3 | Pacing
1.9.6 | Must allow defibrillation with paddle or pad placement as either:
1.9.6.1 | Anterior / Lateral
1.9.7 | Manikin and software must have a feedback mechanism on compressions (rate, depth interruption time), and respirations (excessive ventilation, rate)
1.1 | Must come with a product user guide
1.11 | Must come in a protective case