## Sheet1
Diagnostic Middle ear analyzer (tympanometry and acoustic reflex measurements) 226Hz
Description
A diagnostic immitance unit used to perform tympanometry, ipsi- and contra-lateral reflexes, reflex decay and Eustachian tube function.
Technical Specification
Must have 226Hz probe tone
Must have test reflexes at 250, 500, 1000, 2000 and 4000Hz (pure tone)
Must have 400-4000Hz BBN
Must have clicks in the range of 35-120dB
Must have test tympanometry across range of +200 - -400daPa with accuracy of ±10daPa
Must include a eustachian tube function
Must have a USB PC interface
Must come with the following:
probe assembly (including contra-lateral insert earphone)
probe mount kit/accessory (or shoulder strap)
ear tips (minimum six different sizes, with minimum three of each tip size)
calibration test cavity
cleaning kit
Accessories:
probe assembly (including contra-lateral insert earphone)
probe mount kit/accessory (or shoulder strap)
ear tips (minimum six different sizes, with minimum three of each tip size)
calibration test cavity
cleaning kit
Optional accessories:
eustachian tube function
reflex decay test
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list