## Sheet1
Electrophysiology testing equipment with ABR (air and bone conduction), ASSR, DPOAE
Description
Device used for assessing the peripheral auditory system from the cochlear. Base system which must be pre-loaded with diagnostic ABR software with the ability to add additional modules.
Technical Specification
Complete evoked potential system able to perform diagnostic ABR (click, tone burst, chirp), auditory steady state response (ASSR), diagnostic TEOAES and DPOAEs, middle latency response (MLR), late latency responses (LLR)
Must include laptop, software, clearline adaptor and free field aided testing capacity
Laptop requirements are a minimum of 5GB free disk space, minimum 1GB of RAM, latest Intel version
Base unit must come pre-loaded with ABR software
Must include a minimum of:
2 sets of cup electrodes
2 sets of snap electrodes
pack of electrode pads
OAE nubs
2 packets of large insert earphones (yellow)
2 packets of small insert earphones (pink)
2 packets of infant insert earphones (red)
Circumaural headset
bone conductor
probe cables
carry case
2x Nuprep gel
2 tubs of EEG conduction paste
The service provider will be required to load additional software onto base system once ordered.
Accessories:
TDH Headphones
Bone vibrator
2 sets of cup electrodes
2 sets of snap electrodes
pack of electrode pads
OAE nubs
2 packets of large insert earphones (yellow)
2 packets of small insert earphones (pink)
2 packets of infant insert earphones (red)
Circumaural headset
bone conductor
probe cables
carry case
Consumables
2x Nuprep gel
2 tubs of EEG conduction paste
Optional accessories
Test software for ASSR which includes diagnostic ABR
Test software for TEOAEs and DPOAEs which includes diagnostic ABR
Test software for MLR and LLR which includes diagnostic ABR
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list