## Sheet1
Diagnostic Audiometer (with extended high frequency testing capabilities)
Description
Non PC-based audiometer used to test and diagnose hearing thresholds.
Technical Specifications
Non PC-controlled, clinical two-channel audiometer that must test in increments of no more than 5db
Capable of performing masked and unmasked pure tone and warble tone air and bone conduction and speech testing
Masking sounds must include narrowband, white and speech noise
Frequency range for air conduction must be within a range of 125Hz to 20 000Hz with intensity range from -10dB to 120dB
-10dB to 80 or 100dB SPL Intensity range for air conduction (125-250Hz): -10dB to 100dB decibel range for bone conduction: 250-4000Hz
Must have the option to add following special tests:
SISI
ABLB
stenger
bekesy
weber
tone-in-noise
speech-in-noise
Must come with:
earphones
paired insert phones
bone conductor
response button
talk-forward
built-in or gooseneck microphone or operator's headset with microphone
pair of sound-field speakers
The audiometer offered must allow for user selectable attenuation over the entire range
Must provide auditory as well as visual patient response monitoring from its frequency and intensity controls in order to further enhance user control of its operation
Pre-recorded speech wordlists in English, with the option of South African languages in male and female voice, must be built-in, failing which the audiometer must have a built-in CD player or must have a USB connection and come with a compatible external CD Player
Contralateral white noise, speech noise and narrow band noise for air / bone and free field audiometry masking must be provided on the unit offered
Must be supplied with field speakers
Specifications of field speakers are as follows:
Space saving, high efficiency, flat response, low distortion speakers
Speakers should have a frequency response range of 250Hz to 6000Hz with extendable range up to 8000Hz
Meets ANSI or equivalent standards.
Impedance of 8 Ω (Ohm) and efficiency of 100dB HTL Speech Level
Must be compatible with booths and diagnostic audiometers
Must include mounting brackets for installation
Accessories:
earphones
paired insert phones
bone conductor
response button
talk-forward
built-in or gooseneck microphone or operator's headset with microphone
field speakers in terms of the specifications above
Optional accessories:
SISI
ABLB
stenger
bekesy
weber
tone-in-noise
speech-in-noise
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list