## Sheet1
Screening audiometer and tympanometer (226Hz)
Description
Audiometry equipment used for screening of hearing, tympanometry and acoustic reflexes
Technical Specification
Compact, lightweight and portable screening tympanometer and audiometer combination unit
Must be supplied with a carry case
Must have A/C power supply and battery
Tympanometry measurement capability of 226Hz.
Must have printing capabilities for audiometer and tympanometer
Must be supplied with two rolls of paper
Must have a memory for at least 12 tests, with probe assembly
Must be supplied with tympanometry probes/ear tips
Must be supplied with a box of tympanometry ear tips of at least six different sizes, with at least three tips of each size
Must include ipsilateral reflexes with frequency range from 500Hz to 4000Hz.
Screening audiometer module must include a testable frequency range of 125Hz - 8000Hz and decibel intensity range of -10dBHL to at least 100dBHL
Must be supplied with TDH39 or equivalent head phones
Must be supplied with 1 set of spare audiocups
Must be supplied with patient response button
Accessories:
set of headphones
set of insert earphones
set of audio cups
patient response button
Must be compliant to ANSI S3-6 (Type 4) or equivalent (certificate required)
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list