## Sheet1
Screening audiometer
Description
Portable audiometer for hearing screening
Technical Specification
Compact, lightweight and portable screening audiometer
Must be supplied with protective carry case
Must be supplied with a response button and 1 set of spare audiocups
Is a 1 channel audiometer with testable frequency range of 125Hz - 8000Hz
Must have decibel intensity range of -10dBHL to at least 100dBHL
Must be supplied with TDH39 or equivalent head phones
Must have A/C power supply and battery
Accessories:
a response button
1 set of spare audiocups
Must be compliant to ANSI S3-6 (Type 4) or equivalent (certificate required)
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list