## Sheet1
Screening Automated ABR
Description
Portable Device used for screening of the auditory pathway up until the brainstem.
Technical Specification
The unit must have the following parameters:
stimulus types: Click 90-100 µsec
ABR waveform capture and display graphically
must be able to archives tests
stimulus Rate: 32 to 62 per second
stimulus Intensity: 35dB in 5dB steps
stimulus Polarity: Rarefaction
input frequency Range: 30Hz to 3,000Hz
waveform latency delay: ±0.2 msec
gain: 40-45 dB
Should contain integrated electrodes pads or insert earphones.
Provide inserts of various sizes including neonates
Laptop requirements are a minimum of 5GB free disk space, minimum 1GB of RAM, latest Intel version
Must be able to be connected to a printer for printing purposes if laptop option not available
Accessories:
electrodes pads
insert earphones
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list