## Sheet1
Vestibular equipment: Video Head Impulse Test (Vhit)
Description
Medical equipment used to detect semi-circular canal dysfunction by measuring the patient's vestibule-ocular reflex in response to head movements.
Technical Specifications
Must have VHIT camera
The camera must have the following features:
USB interface of 2.0 or 3.0 to PC
can be powered through USB port
monocular inputs
sample range of 50-300 fps and 30-120Hz
eye tracking of 100-400 pixels, range: horizontal (25-30) and vertical (20-30)
laser: wavelength (max 660nm) and output power (max 0.9mW)
goggle dimensions: length (max 18.4cm), width (1-4.5cm), height (max 4.5cm) and weight (max 60g)
2 built-in calibration lasers for biologic calibration
Must have VHIT software
Must have user manual and installation CD
Must have adjustment features – motorized adjustments, of image sharpness, camera horizontal orientation and camera height
Must have weight of 6kgs (without cable)
Must have USB connection to PC
Camera must have definition of 752 x 480; cell type of CCD 1/3; power supply (12 v DC); sensitivity of < 0.005 lux, and a sampling rate up to 100Hz
Must have infrared lighting (wavelength of 830nm and light cone of +/- 22 deg.par LED)
Must have power supply (input of 100-240 VAC/ 50-60Hz/ Max 0.9A; output of 13.2 V DC – 3A; max. output power of 40W; class II insulation, and an operating altitude of <2000m)
The VHIT must be CE approved/ EN 60601-1/ EN 60601-1-2/ IEC 62471-1
Minimum computer requirements:
Operating system: 64 bit Windows 10
CPU: Intel i7 processor
Memory: 16GB
Disk space: 500G
Connectors: up to four USB 3.0 connectors
Graphics card: card with min 2G memory
Monitor: High resolution
Components: mouse and keyboard, wireless option is supported
Internet access: recommended to be connected upon installation
Must have 3 modes: live, acquisition and results
Must have computer controlled camera settings, on 3 axes (horizontal and vertical alignment of the patient in image and sharpness)
Must have adjustment feature: interpupillary distance, acceleration threshold
Must have protocol examination guide (visual indicators, synthetic voice messages)
Must have auto detection of head feature
Must have manual correction available
Must have VOR high-frequency analysis
Must have feature that allows detection and analysis of saccade strategies (including overt and covert saccades)
Software must have graph feature (right/left gaze, head movement, position, velocity and gain)
Must have playback feature - playback of video recording (slow motion showing the deviation of the gaze, synchronised with graphs)
Must comply with IEC safety standards
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list