## Sheet1
Visual Response Audiometry unit (Video) with DVD player
Description
Equipment used together with audiometer for paediatric hearing evaluations consisting of toy cabinets or video re-enforcers and sound-field speakers.
Technical Specifications
A Video VRA reinforce unit with DVD player
Two colour LED flat screen 15" - 20" monitors that display DVD recording for the use during paediatric hearing testing.
Switch must be able to be installed in the observation room to control animation (i.e. on and off), and switch on the screen.
No time elapse should occur between the initiation of VRA system and illumination and activation of toy reinforce/ screen.
Unit must come with DVD player option to provide video stimulus. Must include x2 mounting brackets for installation of the screens.
Each screen should be able to be lit up independently.
2 colour monitors or 2 boxes with toys that light up and move, compatible with diagnostic audiometer and intermediary compute/laptop system with HDMI connecting cables to be included if system requires
Reinforcement activated by push of button by either wireless remote or through computer system
Silent switch that can be connected to the audiometer for easy changes between right and left monitors
Variety of animated characters in reinforcement playlist
Must include x2 mounting bracket for installation to mount flush against wall or on arm lever
Must be compatible with audiometer installed and sound field speakers if not included
Must include AC power supplies
Must include switch for easy changes between right and left monitors
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list