## Sheet1
Visual Response Audiometry unit (4 toys x 2 sides)
Description
Equipment used together with audiometer for paediatric hearing evaluations consisting of toy cabinets or video re-inforcers and sound-field speakers.
Technical Specifications
The unit must be a 4 toys x 2 sides i.e. 2 cabinets containing 4 toys each
lluminated and animated child friendly toys, placed within a cabinet, obscured by smoked perspex screen
Switch must be able to be installed in the observation room to control animation (i.e. on and off), and light up/ illuminated the toys
Each toy should be able to be lit up independently
No time elapse should occur between the initiation of VRA system and illumination and activation of toy reinforcer
Must include x2 mounting bracket for installation of the cabinets
2 colour monitors or 2 boxes with toys that light up and move, compatible with diagnostic audiometer and intermediary compute/laptop system with HDMI connecting cables to be included if system requires.
Reinforcement activated by push of button by either wireless remote or through computer system.
Silent switch that can be connected to the audiometer for easy changes between right and left monitors
Variety of animated characters in reinforcement playlist
Must include x2 mounting bracket for installation to mount flush against wall or on arm lever
Must be compatible with audiometer installed and sound field speakers if not included
Must include AC power supplies
Must include switch for easy changes between right and left monitors
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list