## Sheet1
Fixed Socket Router Machine (complete)
Description
A fixed machine used for grinding and milling of orthotics and prosthetic materials
Technical Specification
The unit must have 5/8" milling shaft thread with adapter from 5/8" to M16 thread adapter 5/8"
The unit must be a 3 phase heavy duty device
The unit must consists of suction hood/extraction shroud with swivelling mounting rods
The dust collection/suction system should be compatible with ordering facility
The unit must have a shaft protector short and long
The unit must have a exhaust nozzle for external suction system
The unit must be fitted with adjustable mount for the extraction shroud and flexible hose with diameter of 125 mm
The tool set for changing the socket router cutters, exhaust arm with clamping lever for positioning , tool holder,
The unit must have a main safety switch
The unit must have a on/off button for router motor gear
The unit must have a emergency stop foot switch
The unit must have a emergency stop button
The unit approximate dimensions (WxDxH) 1,150 x 930 x 1,650 mm
The exhaust nozzle should be approximately 125 mm
A minimum of 300mm electric height adjustment
Variable rotation speed, min. 1 500 RPM
Two speeds and electro magnetic brake.
Router shaft length should be a minimum of 450 mm
Connection to a central chippings extraction system with a minimum air velocity of 20 m/s ( exhaust nozzle diameter 125mm)
The unit should include:
open end spanner/wrench- 27 mm
chip/dust extraction shroud - complete
thread stud - for tools with 5/8" thread
threaded stud - for tools with M16 thread
1 mounting/fastening set
thread protector nut - 5/8" inside thread
shaft guard - long, shaft guard short
floor fastening set
Accessories :
27 mm wrench
floor fastening set
threaded stud
protective nut
extraction shroud
protective shaft housing, long
protective shaft housing, short
The unit must have an approximately 2m power cord with CEE plug and CEE socket
Electrical connection in (V/Hz/kw) 3x400N/Pe / 50-60 / 1.5
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list