## Sheet1
Normal Socket Router Machine
Description
A machine used for grinding and milling of orthotics and prosthetic materials
Technical Specification
For connection to a central chD57:D87ipping/suction exhaust system
Min air velocity of 20 m/s and negative pressure of 703 Pa
Exhaust connection pieceǾ125mm
Two adjustable speeds
Electro-magnetic brake
Extraction shroud with swivelling mount and flexible hose
Removable protective shaft housings, long and short lengths
Height between 1.350 - 1.500mm
Shaft length not less than 450mm
Not weight more than 250kg
Accessories:
27mm spanner/wrench
floor fastening set
threaded stud 5/8
protective nut
extraction shroud/ suction hood
all the swivelling mounts
flexible hoses
a short and long protective shaft housing
Electrical requirements:
CEE plug and socket 5 x 16 A
power cord not less than 2 m
emergency power off foot switch
2 speed selection switch
motor protection switch
V/Hz/kW - 3 x 400 N/PE /50-60/ 1.5
1.500/3.000 RPM min¯¹
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list