## Sheet1
Adjustable Adult electric standing frame with rechargable battery back up
Description
Rehabilitation device to assist patients with poor standing tolerance
Technical Specification
Standing Frame: Suitable for patients from 105 to 200cm in height.
Contoured recess table, which is height adjustable.
footboard with contoured recess for wheelchairs.
Easily movable (with wheels/castors) and brakes (lockable castors).
Easy to clean and sterilise.
Made of highly durable steel frame, with non-slip leather covering and suitable padding.
Approximate dimensions: Table adjustable in height in tilt (0-23°) and backwards or forwards.
Electronic lifting band (optional Maxi size) with internal covering in non-sliding shammy leather.
Hooks to allow traction while climbing.
Anterior and posterior trunk support with a minimum of two straps:
Adjustable knee height with knee padding and knee support (adjustable in height, width and backwards/forwards).
Adjustable heels guards (adjustable in width and backwards-forwards)
Distance from arm support to ground 90-110 cm
Distance from table to footrest cm ±85min to ±125max
Minimum width of frame 80cm
Total length ±100-115
Dimensions of the table 65-75 x 55-65 cm
Standard lifting band width x height cm 85 x 28
Alternative maxi size lifting band width x height cm 100 x 30
Alternative lifting band with pelvis belt width x height cm 85/123 x 39
Alternative maxi lifting band with pelvis belt width x height cm 100/138 x 41
Load: 150 -200kg
Rechargeable batteries that last for 1 hour when fully charged
Accessories:
Chest and back support
pelvis support
front foot restraints with belts
lifting band with pelvis belt (maxi size included)
Rechargeable batteries
Batteries charger
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list