## Sheet1
Adult Electric Tilt Table
Description
Electrically operated tilt table has been designed to assist in the therapy of juvenile or adult clients who are unable to stand on their own.
Technical Specification
Tilt table, Adult: Electrical, Variable height.
Minimum of 3 Straps to secure patient.
Overall tilt 0° -90°.
Must have tilt locking mechanism.
Comfortable foot rest with non-slip surface.
With manual override back up option in case of emergency
Load capacity: Minimum 200kg
Optional Wheel raising mechanism for easy movement of the bed or minimum 4 wheels each with brakes.
Height High Position ±90cm, Low Position ±45cm (horizontal)
Overall Length: not less than ±180cm
Overall Width: not less than ±60cm.
Optional tapering of head and foot piece to a minimum of 50cm.
FOAM DENSITY: Standard: 38 Kg/m3 and 4-6cm thick.
Upholstered surface with room for both patient and therapist imitation leather upholstery contains no phthalates and is flame retardant.
Material must be water proofed, resistant and disinfectable
Accessories:
work surface
set of 3 straps
Optional feature:
Rechargeable back-up battery
Battery charger
side handles
Minimum 3 metre electrical cord.
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list