## Sheet1
Hand evaluation kit
Description
A set of instruments used to test grip strength and other components of hand function
Technical Specification
Kit includes: hydraulic hand dynamometer, pinch gauge, finger goniometer, Sensory evaluator hand kit, Two point Discriminator, Tape measure, finger circumference gauge (cm reading).
Hydraulic hand dynamometer:
Maximum strength indicator remains after each reading until reset.
Registers in pounds and kilograms.
Approximate Dimensions: 28cm x 10cm x 20cm
Reliable grip force evaluations with adjustable handles which can be placed in 5 grip positions from 35mm to 87mm in 13mm increments to accommodate the hand size.
To include calibration certificate and norm ranges.
Hydraulic pinch Gauge:
It offers accurate and repeatable pinch strength readings (tip, key and palmer).
Maximum strength indicator remains after each reading until reset.
Registers in pounds and kilograms.
Approximate Dimensions: 13cm x 5cm x 6cm
To include calibration certificate and norm ranges.
Finger goniometer:
6’ stainless steel finger goniometer
Has 180 degrees scale marked in 1 degree increments that read in either direction.
Linear in inches and centimeters.
Must be able to measure hyperextension of the finger. (0 - 120° with hyperextension 0 - 30°)
Sensory evaluator hand kit
5 piece sensory evaluator kit
Each evaluator is individually calibrated within a 5% standard deviation of a predetermined level of force
The sensory evaluator handles are marked with the corresponding color-coded pressure level
The following colour coded pressure levels must be included: 2.83, 3.61, 4.31, 4.56, 6.65
A small carry case muse be included for storage
Two point Discriminator
Consists of two sturdy, rotating, plastic disks that are joined together
Rounded tips are spaced at standard testing intervals from 1 to 15mm apart
A 20 and 25mm spacing are also given
One disk setting tests from 1 to 8mm. To change to the 9 to 15mm disk setting, simply rotate the top disk until it clicks in place.
Instructions included
Tape Measure
Small, compact measuring tape used to measure the circumference of the digits and hand
Measures in inches and centimetres
Must be flexible to wrap around the digit/ hand
Must be retractible
Measures up to a minimum of 1 metre
Finger circumference gauge reading
Used to measure the circumference of a finger
Finger circumference gauge webbing loop wraps around the finger and the circumference is read off the ruler
Must be measured in inches and centimetres
Accessories
Hard carry case for storage of the 7 piece kit
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list