## Sheet1
Dual Channel TENS (Trans Electrical Neuromuscular Stimulation) and NMS (Neuromuscular Stimulation) Unit
Description
A device used for pain relieve and muscle rehabilitation
TECHNICAL SPECIFICATION
The unit shall comply with at least the following specifications:
Waveform: Asymmetrical, rectangular, biphasic with zero DC current
Max Voltage: 180Volts +10/-30V
Pulse Amplitude: Adjustable from 0-90mA
Pulse Width: Adjustable from 50 - 450µS(microseconds)
Pulse Rate: Adjustable from 2 - 200Hz
Lock mode function
The unit shall have the following features:
At least 2 totally independent channels
Possible to run to different programs at the same time
A minimum of 10 present TENS programs
A minimum of 2 present EMS programs
A minimum of 2 customisable programs
A large digital display
Clear backlit LCD display
Rechargeable Battery operated for ease of portability
Must be supplied with a battery charger
Accessories:
Re-usable electrodes in different sizes
Disposable electrodes in different sizes
Dual conductor lead wire for TENS/NMS
Optional accessories
Rechargeable battery
Battery charger
The unit shall be FDS or CE approved (Attach certification)
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list