## Sheet1
Electronic Height Adjustable Plinth 5 Section
Description
An electrical operated 5 section unit used to position patients during physical rehabilitation procedure
Technical Specification
The unit shall have an electric hi-low height adjustment with high position ≥90cm and low position ≤45cm
The unit shall have a remote control and foot control unit
The unit shall have a load capacity of 200kg+
Overall length shall be a minimum of 180cm
Overall width shall be a minimum of 60cm (Optional tapering of head and foot piece to a minimum of 50cm)
The frame shall be high grade and durable
The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists
All adjustment handles shall be durable and made of stainless steel or equivalent quality material
The unit shall have heavy duty castors with brakes for easy movement of the unit
The castors shall be fully retractable
The feet that the unit stands on when the castors are retracted, shall allow for individual adjustment to allow levelling on uneven floors
The feet shall be heavy duty
The unit shall be upholstered
Upholstery foam density shall be at least 38 Kg/m3 and 4-6cm thick.
The upholstery material shall be durable, hygienic ,non slip and washable (e.g. imitation leather)
The upholstery shall be free of phthalates and shall be flame retardant
The unit shall comprise of five sections that are adjustable at four points (head and lumbar and two arm rests)
The unit shall accommodate lying and sitting position,
The head section shall have a comfortable breathing slot with a plug
The head section shall have backrest and angle shall allow tilt from at least -30° to +90°
Accessories
replacement castors
The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS
The unit shall be FDS or CE approved (Attach certification)
The unit shall be supplied with operators and service & maintenance manuals
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list
HYDRAULICS