## Sheet1
Neurology Plinth-Hydraulic
Description
A wide hydraulic operated unit used to position patients during active physical rehabilitation procedure
Technical Specification
The unit shall have a hydraulic lifting system that is manually operated
Height adjustment with high position up to 90cm and low position to ≤45cm
The unit shall be able to lift and carry a load of 250kg+
Surface area should be equal or better than 120cm width and minimum of 180cm length
The frame shall be high grade and durable
The frame shall have a ground clearance of at least 14cm to allow for use of patient transfers units such as hoists
All adjustment handles shall be durable and made of stainless steel or equivalent quality material
The unit shall have heavy duty castors for easy movement of the unit
The castors shall be fully retractable
The feet that the unit stands on when the castors are retracted, shall allow for individual adjustment to allow levelling on uneven floors
The feet shall be heavy duty
The unit shall be upholstered
Upholstery foam density shall be at least 38 Kg/m3 and 4-6cm thick.
The upholstery material shall be durable, hygienic and washable (e.g. imitation leather)
The upholstery shall be free of phthalates and shall be flame retardant
The unit shall comprise of two sections that are adjustable to accommodate a lying and a sitting position
Accessories
replacement castors
The unit shall be FDS or CE approved (Attach certification)
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list