## Sheet1
Stand-Up Manual Patient Lift
Description
A stand-up unit used for patient transfer
Technical Specification
Non-powered (non electric)
Safe minimum loading capacity of 180kg
Weigh no more than 30kg
Length x height: ±900mm x ± 1050mm
Turning radius: ±1150mm
Minimum door requirement 630mm
Pedal operating force <300N
Pedal operated adjustment of leg width
Internal width legs open ±730mm
Low friction castors
Compliance with legislature and applicable product standards
Shall have the following features:
Manoeuvering handles
Cross bar/handle bar
Sling attachment lugs/clips for back support strap
Pivoting seat halves/seat pads
Knee pad
2 Front castors
2 Back castors with brakes
Footboard/base
Leg pivot foot levers/pedals
Expected frame life of at least 10 years
Expected operational life of slings/backrest of at least 2 years
Accessories:
sling/back rests in a variety of sizes/adjustable size
sling/back rests to be manufactured of a durable material
sling/back rests shall be compatible with the purchased stand-up manual patient lift
secure buckle for fastening around patient
secure clips for fastening to frame
sling attachment lugs/clips for back support strap
knee pad
castors
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list