## Sheet1
Autoclave
Description
Technical Specification
Mini portable bucket like unit
The unit must come with 3 Instrument Cases/ cassetes, or Frame/ holder to hold the cases / cassette during sterilization and tongs to remove instrument case or cassette after sterilization
Fully automatic pressure for the effective sterilization of hand pieces and instruments sterilizable in autoclave.
It is made of aluminium alloy and has 5 safety devices
For Steam Sterilization/disinfection of small surgical instruments
With Handles for Lid and Base.
Handle to allow easy transportation with one hand
Size 9 Litres capacity.
Support for 6 cases/ cassets
Works at 121º C and maintain this temperature for 11 minutes together with the heating time and the cooling phase,
The total sterilization cycle lasts approximately 25 minutes
It is equipped with an internal support for 6 boxes where the instruments are placed.
Features:
light, compact and portable
single, one step operation
fully automatic
large capacity
safe and reliable
up to 6 cassettes for large capacity
advanced design
a simple, one step operation starts the cycle
large sequence indicators show: "Power unit" - "Cycle in Progress" - "Sterilization in Process" - "Cycle is Successful" or "Cycle has Failed"
an interlock system prevents the lid from being removed while pressure remains in the vessel
a two stage "over pressure" protection system incorporates a calibrated Pressure Release Valve and a Gasket Offset Device
electronic detectors turn off the power if there is insufficient water
optional temperature and pressure gauges show internal chamber conditions during the cycle
TST indicator strips provide independent verification that the correct combination of temperature, steam and time has been achieved for successful sterilization
Temperature: 121° C (250° F)
Sterilization Time: 16 minutes
Cycle Time: Approx. 30 minutes
Accessories:
additional lid gasket
case/cassets- instrument cassets/ cases plastic additional 3 cassetes
2 m Power cord
Electrically operated, Power source 220v/50hz
For your assurance of quality, each autoclave is computer tested and issued with a unique Performance Test Certificate showing that the parameters for successful sterilization have been met.
All accessories and consumables must be supplied standard with the unit
A separate price of each accessory and consumables listed above must also be supplied with this offer
Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list