Improving quality of Life and Enhancing Sustainable Economic Development

SPECIFICATIONS OF GOODS REQUIRED

- Background

UMngeni–uThukela Water requires provision of advanced condition monitoring tool for South Central Region.

South Central Region assets consist of numerous MV/LV switchgears rotating which operates 24/7 and require advanced condition monitoring for effective functionality. Therefore, this advanced conditioning tool shall assist in conducting predictive (condition-based) maintenance where certain parameters will be monitored to detect signs of imminent.

Performing advanced condition assessment of medium voltage (MV) switchgear is essential for ensure the safety, reliability, efficient operation and compliance with SANS 10142 and SANS 62271 standards. This process involves detailed inspection, diagnostic testing and performance evaluation to determine the health and remaining life of the switchgear of electrical distribution systems.

  - Condition Monitoring Tool Technical Specifications

- OMICRON CIBANO 500

| Power output of integrated power supply |  |  |  |
| --- | --- | --- | --- |
| Frequency | DC / 15 Hz ... 400 Hz |  |  |
| Power | Vmains | P30s | P2h |
|  | - 100 V | 1500 W | 1000 W |
|  | - 190 V | 3200 W | 2400 W |
| Current / voltage output of integrated power supply |  |  |  |
| Source | Range | Imax, 30 s1 | Imax, 2 h1 |
| DC | 0 ... ±300 V | 27.5 A | 12 A |
| DC | 0 ... ±150 V | 55 A | 24 A |
| AC | 0 ... 240 V | 20 A | 12 A |
| AC | 0 ... 120 V | 40 A | 24 A |
| Commands for control of trip or close coils |  |  |  |
| Current per channel5 |  | Duty cycle |  |
| 6 Aeff AC or DC |  | Continuous |  |
| 15 Aeff AC or DC |  | 20 s on |  |
| 30 Aeff AC or DC |  | 80 s off |  |
| 40 Aeff AC or 55 A DC |  | 10 s on |  |
|  |  | 190 s off |  |
|  |  | 200 ms |  |
| Commands for motor supply |  |  |  |
| Current per channel5 |  | Duty cycle |  |
| 24 Aeff AC or DC |  | Continuous |  |
| 40 Aeff AC or DC |  | 20 s on |  |
|  |  | 80 s off |  |
| 55 A DC |  | 10 s on |  |
|  |  | 190 s off |  |
| Voltage measurements |  |  |  |
| Source |  | Range | Accuracy |
| DC |  | 0 ... 300 V | 0.1 % rd + 0.05 % fs |
| AC |  | 0 ... 300 V | 0.03 % rd + 0.01 % fs |
| DC |  | 0 ... 3 V | 0.1 % rd + 0. 05 % fs |
| DC |  | 0 ... 300 mV | 0.1 % rd + 0.1 % fs |
| DC |  | 0 ... 30 mV | 0.1 % rd + 0.1 % fs |
| Current measurements |  |  |  |
| Source |  | Range | Accuracy |
| DC |  | 0 ... 55 A | 0.1 % rd + 0.2 % fs |
| AC |  | 0 ... 40 A | 0.1 % rd + 0.1 % fs |
| Mains supply |  |  |  |
| Voltage |  | Nominal: 100 V ... 240 V AC |  |
|  |  | Permitted: 85 V ... 264 V AC |  |
| Current |  | Nominal: 16 A |  |
| Frequency |  | Nominal: 50 Hz / 60 Hz |  |
|  |  | Permitted: 45 Hz ... 65 Hz |  |
| Power fuse |  | Automatic circuit breaker with magnetic |  |
|  |  | overcurrent tripping at I > 16 A |  |
| Power consumption |  | Continuous |  |
|  |  | < 3.5 kW Peak |  |
|  |  | < 5.0 kW |  |
| Voltage input from station battery |  |  |  |
| Source |  | Range | Accuracy |
| AC |  | 0 ... 300 V | 0.5 % rd + 0.5 % fs |
| Resistance measurements |  |  |  |
| Range | Voltage range | Injected current | Accuracy |
| 0.1 μΩ ... 300 µΩ | 30 mV | 100 A | 0.2 % rd + 0.1 μΩ |
| 0.5μΩ...3 mΩ | 300 mV | 100 A | 0.2 % rd + 0.5 μΩ |
| 5 μΩ ... 30 mΩ | 3 V | 100 A | 0.2 % rd + 5 μΩ |
| 50 μΩ ... 300 mΩ | 3 V | 10 A | 0.2 % rd + 50 μΩ |
| Mechanical data |  |  |  |
| Dimensions (W × H × D) |  | 580 × 386 × 229 mm |  |
| (W = 464 mm without handles) |  |  |  |
| Weight |  | 20 kg (main unit including power supply) |  |
| Environmental conditions |  |  |  |
| Temperature |  | Operating: -10 °C ... +55 °C Storage: -30 °C ... +70 °C |  |
| Relative humidity |  | 5 % ... 95 %, non-condensing |  |
| Maximum altitude |  | Operating: 2 000m up to 5000 m with limited specifications, according to footnotes 2 and 4 |  |
| Storage |  | 12000 m |  |
| Equipment reliability |  |  |  |
| Shock |  | IEC / EN 60068-2-27, 15 g / 11 ms, half-sinusoid, 3 shocks in each axis |  |
| Vibration |  | IEC / EN 60068-2-6, frequency range from 10 Hz to 150 Hz, continuous acceleration 2 g (20 m/s2 ), 20 cycles per axis |  |
| Inputs for auxiliary contacts |  |  |  |
| Auxiliary input type |  | Toggling with potential-free (dry) contacts or voltages (wet) up to 300 V DC |  |
| Maximum sample rate |  | 40 kHz |  |
| Minimum resolution |  | 25 μs |  |
| Interfaces |  |  |  |
| Digital |  | 1 × Ethernet, 1 × Serial, 2 × Safety |  |
|  |  | Optional Ether-CAT module: 4 × Ether-CAT Optional auxiliary module: 1 × Ether-CAT |  |
| Analog |  | 1 × analog input (V IN) |  |
|  |  | 3 × analog input/ analog output/ binary input (A) |  |
|  |  | 4 × analog input / analog output (B) |  |
|  |  | Optional auxiliary module: 3 × binary input (C) |  |
| PC Requirements |  |  |  |
| Operating system |  | Windows 10 64-bit Windows 8.1 64-bit Windows 8 64-bit Windows 7 SP1 64-bit |  |
| CPU |  | Multicore system with 2 GHz Single-core system with 2 GHz |  |
| RAM |  | 4 GB |  |
| Hard disk |  | 5 GB of available space |  |
| Storage device |  | DVD-ROM drive |  |
| Graphics adapter |  | Super VGA (1280 × 768) or higher-resolution video adapter and monitor |  |
| Interface |  | Ethernet NIC / USB 2.0 |  |
| Microsoft software |  | Microsoft Office 2016, 2013, 2010, or 2007 |  |

OMICRON CIBANO 500 standard package consisting of the following:

- CIBANO 500 test hardware

- PTM Advanced software

- Test functions for static contact resistance measurement, main and auxiliary contact timing, analysis of coil currents

- Hardcover transport case with wheels

- Manual, power lead, PC connection lead, standard measurement leads, combination lead for A outputs and B outputs, connection accessories, grounding lead