{"text_format":"3","workbook":"xlsx","n_sheets":3,"sheets":["1. Laser Co2","2. All TISSIUE","Sheet1"],"not_rendered":{"images":2,"defined_names":2},"read":"whole","date_system":"1900"}

{"sheet":"1. Laser Co2","index":0,"n_rows":66,"n_cols":6,"rows_kept":66,"merged":[[0,0,0,5],[1,0,1,5],[2,0,2,5],[3,0,3,5],[5,0,5,5],[6,0,6,5],[7,0,7,5],[8,0,8,5],[9,0,9,5],[10,0,10,5],[11,0,11,5],[12,0,12,5],[13,0,13,5],[14,0,14,5],[15,0,15,5],[16,0,16,5],[17,0,17,5]]}
["ANNEXURE A. : TECHNICAL SPECIFICATION"]
["ITEM 1"]
["GT/GDH/054/2023:THE SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF LASER Co2 MACHINE TO VARIOUS GAUTENG HEALTH INSTITUTIONS FOR A PERIOD OF THREE YEARS"]
["TECHNICAL SPECIFICATION FOR LASER Co2 MACHINE"]
["Item no.","Technical Specification","Weight","Comply (Yes/No)","Details of Product Offer                                            (Provide your answers in this Colum. You are advised to be straight to the point)","Reference to Brochure / Technical Data Sheet (verification) /  Supporting Document"]
["Note:                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                   "]
["1. Technical Compliance"]
["(a) All Equipment tendered must comply with all the specifications with a # weight\n(b)All equipment tendered for must score 3 or higher technically compliant to qualify for consideration for an award"]
["2. Scoring of the bids"]
["2a. Weight: Each specification will have a weight of 1 or a hash (#), which indicates the level of importance, as follows:   "]
["1 -  A specification with a standard level of importance that must be complied with.,"]
["# -  Indicates a specification with a essential level of importance, where the bid offer will be disqualified, if the specification is not complied with."]
["2b. Score: The Bid Evaluation Committee will assign a score of 0 or1 for each specification, which indicates the level of compliance, as follows:"]
["0 -  The specification does not  comply.     1 -  The specification is fully compliant."]
["2c. Weighted and Total Score: The score will be multiplied by the weight to give the total score for each specification and be added up. "]
["2d. A score of 3 or higher compliance: In order to be considered for an award, the bidders must comply as follows with the specification: "]
["i. The bidders must comply with all the specifications (line items) with a # weight. "]
["ii. The bidders must obtain a minimum score of 3 or higher out of the total score of 5 points in respect of specifications (line) with a weight of 1."]
["A.","LASER Co2 MACHINE","Weight","Comply (Yes/No)","Details of Bid Offer                                            (Provide your answers in this Colum. You are advised to be straight to the point)","Reference to Brochure / Technical Data Sheet (verification) /  Supporting Document"]
["A.1","It must be a medical carbon dioxide laser with a wavelength 10600 micro meters, far infrared.","#"]
["A.2","The maximum must not be less than 60 watts power, and the power must be adjustable from 1 watt to a maximum.","#"]
["A.3","It must have minimum of 3mw red diode aiming beam, 635nm, adjustable intensity.","#"]
["A.4","It must be microprocessor based.","#"]
["A.5","it must have a radio frequency (RF) sealed co2 or direct current (DC) excited laser source.","#"]
["A.6","It must have continuous, single pulse and repeat pulse tissue exposure modes.\n","#"]
["A.7","it must have an average continuous power of 01 Up to 60 watts.","#"]
["A.8","the beam delivery must be through a joint, counter Weight or spring balanced arm, with 360 deg rotation.","#"]
["A.9","The laser arm must be long enough not to compromise the microscope's movement and not to put strain on the laser arm.","#"]
["A.10","Articulated arm must be protected during storage. Please indicate how this is done.","#"]
["A.11","It Must Have A Timed Exposure Of Following Durations:"]
["A.11.1","On time (single pulse) between 0.01 – 1 . 00 sec.","#"]
["A.11.2","On time (repeat pulse) between 0.01 – 1.00 sec.","#"]
["A.11.3","it must have a repeat delay, off time, 0.1 to 5.0 sec","#"]
["A.11.4","it should have at least 100 user defined memory settings",1]
["A.11.5","it must have the possibility of saving a limited number of favourite user setting with fast recall","#"]
["A.11.6","it must include handpiece sets x2 and include all other handpiece sets","#"]
["A.11.7","it must include a complete co2 endoscopic endonasal surgery set","#"]
["A.11.8","a bacterial limiting smoke evacuator must be included in the offer","#"]
["A.11.9","it must come with a minimum of ten laser safety glasses","#"]
["A.11.10","it must have the function of equipping scanners with preset recommendations for parameters for different applications","#"]
["A.11.11","it must have a lcd multi –colour high resolution touch screen control panel up to 10.4\" inches",1]
["A.11.12","it must have a graphic display that provide in settings step by step operating instructions","#"]
["A.11.13","it must have a self-contained closed loop cooling system","#"]
["B","Accessories:"]
["B.1","micromanipulator for microlaryngeal laser surgery with the following features:"]
["B.1.1","the system must be able to fully reduce thermal damage to surrounding tissue to promote quicker healing\n","#"]
["B.1.2","it must have an optical design to assure perfect coincidence of the aiming beam diode and co2 beams even at highest microsurgical magnifications\n","#"]
["B.1.3","it must be easily adjustable and must have variable working distance from 200mm to 400mm",1]
["B.1.4","it must have continuously variable defocus with a user adjustable defocus limiter","#"]
["B.1.5","its joystick must be able to rotate scanning shapes, modify their dimension, and having the possibility of Passing from scan-on/scan-off function, without surgeons needing to take their eyes off the microscope oculars","#"]
["B.1.6","it must be user selectable for left and right hand controls","#"]
["B.1.7","it should have a minimum spot size of 250 microns at 400mm effective focus length (efl).",1]
["B.2","The machine must have a scanning “robotic” laser Microsurgery system to reduce operating time, Simplify procedures and enable uniform incision or ablation in a single laser scan, with following Requirements:"]
["B.2.1","it must have different beam scan shapes: i.e"]
["B.2.1.1","A). Linear & curved incisions for cutting","#"]
["B.2.1.2","B). Hexagon and circle for ablation: 0.2mm to 6.3 mm in max scanning dimension (user defined). ","#"]
["B.2.1.3","C). Circular shape dedicated to one shot laser d)stapedotomy: dimensions from 0,3mm to 1,1mm (user defined) with database parameters","#"]
["B.2.2","it must have a penetration depth of 0.2 mm up to 2 mm. (user defined).",1]
["C","The offer must include the following:"]
["C.1","built in second laser source up to 50 watt (diode 980nm or 940nm wavelength) with electronically Control switching between diode fibre and free beam\nWith led illuminated indication of active part\n","#"]
["C.2","fibres must be available in different diameters: I.e. 300, 400, 500, 600 micron sterile, single / multiple use","#"]
["C.3","rigid hand piece kit with bendable autoclavable Cannulas","#"]
["C.4","stripping and cutting tools to reuse fiber","#"]
["C.5","database parameters for diode one shot Stapedotomy\n","#"]
["C.6","the laser system must include a co2 fibre beam delivery","#"]
["D.1 ","TOTAL WEIGHTED SCORE",5]

{"sheet":"2. All TISSIUE","index":1,"n_rows":69,"n_cols":6,"rows_kept":69,"merged":[[0,0,0,5],[1,0,1,5],[2,0,2,5],[3,0,3,5],[5,0,5,5],[6,0,6,5],[7,0,7,5],[8,0,8,5],[9,0,9,5],[10,0,10,5],[11,0,11,5],[12,0,12,5],[13,0,13,5],[14,0,14,5],[15,0,15,5],[16,0,16,5],[17,0,17,5]]}
["ANNEXURE A.: TECHNICAL SPECIFICATION"]
["ITEM 2"]
["GT/GDH/054/2023:THE SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF ALL TISSUE LASER MACHINE TO VARIOUS GAUTENG HEALTH INSTITUTIONS FOR A PERIOD OF THREE YEARS"]
["TECHNICAL SPECIFICATION FOR ALL TISSUE LASER "]
["Item no.","Technical Specification","Weight","Comply (Yes/No)","Details of Product Offer                                            (Provide your answers in this Colum. You are advised to be straight to the point)","Reference to Brochure / Technical Data Sheet (verification) /  Supporting Document"]
["Note:                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                                   "]
["1. Technical Compliance"]
["(a) All Equipment tendered must comply with all the specifications with a # weight\n(b)All equipment tendered for must score 28 or higher technically compliant to qualify for consideration for an award"]
["2. Scoring of the bids"]
["2a. Weight: Each specification will have a weight of 1 or a hash (#), which indicates the level of importance, as follows:   "]
["1 -  A specification with a standard level of importance that must be complied with.,"]
["# -  Indicates a specification with a essential level of importance, where the bid offer will be disqualified, if the specification is not complied with."]
["2b. Score: The Bid Evaluation Committee will assign a score of 0 or1 for each specification, which indicates the level of compliance, as follows:"]
["0 -  The specification does not  comply.     1 -  The specification is fully compliant."]
["2c. Weighted and Total Score: The score will be multiplied by the weight to give the total score for each specification and be added up. "]
["2d. A score of 28 or higher compliance: In order to be considered for an award, the bidders must comply as follows with the specification: "]
["i. The bidders must comply with all the specifications (line items) with a # weight. "]
["ii. The bidders must obtain a minimum score of 28 or higher out of the total scoreof 32 points in respect of specifications (line) with a weight of 1."]
["A.","ALL TISSUE LASER","Weight","Comply (Yes/No)","Details of Bid Offer                                            (Provide your answers in this Colum. You are advised to be straight to the point)","Reference to Brochure / Technical Data Sheet (verification) /  Supporting Document"]
["A.1","All tisue laser for disection of all types of tissues during Maxillofacial oral surgical procedures.","#"]
["A.2","Consumables must be available after implementation of the systems for alteast 10 years or the lifespan of the machine","#"]
["A.3","Components:"]
["A.3.1","All tissue: A unit with components for diiferent tissue cutting specifications (Bone, dental hard tissue and soft tissue)","#"]
["A.3.2","Measuring and analysis software including power data tools, calibration programs, and report generator must be included","#"]
["A.3.3","Transducer type piezo resistive; range 5-100Hz","#"]
["A.4","Performance features:"]
["A.4.1","State of the art Bluetooth technology",1]
["A.4.2","Comprehensive standard equipment: (laser unit, display covers, trunk fiber and fiber support, yellow air tube, protective laser eyeware x3, handpieces x2 and fractional handpiece, fiber tip starter kit(all inclusive tips), tip holder, tip cleaning kit, water filter unit, power cord, footswitch, laser warning signs x3.","#"]
["A.4.3","The all-inclusive test program will last no longer than one or two minutes","#"]
["A.5","Dimensions:"]
["A.5.1","Trolley with swivel mounted desktop screen",1]
["B","Calibration:"]
["B.1","Automatic / manual or machine set calibrations",1]
["B.2","Measurement programs",1]
["B.3","Main control: circuit breaker",1]
["B.4","on/off control : Key switch",1]
["B.5","Water type : Distilled or de-ionised only",1]
["B.6","External Air source : 80-120psi (5.5-8.2 bar)",1]
["B.7","Interaction zone :0.5 - 5.0mm from hadpiece tip to target",1]
["B.8","Laser classification: IV (4)µ","#"]
["B.9","Medium : 2.780nm, Er,Cr:YSGG","#"]
["B.10","Frequency : 5-100 hz",1]
["B.11","Average power : .01 -10.0 W",1]
["B.12","Pulse energy: 0 -600mJ",1]
["B.13","Pulse duration for \"H\" mode: 60 µs",1]
["B.14","Pulse duration for \"S\" Mode: 700µs","#"]
["B.15","Mode: Multimode","#"]
["B.16","Tip diameter : 200 - 1200 µm",1]
["B.17","Administrative Programs/Software",1]
["B.18","Data Base:"]
["B.18.1","Data base tools",1]
["B.18.2","Screen and Printer report",1]
["B.18.3","Interpretation",1]
["B.18.4","Off-line entry",1]
["B.18.5","On-line Data transfer ( e.g. Hz and tissue type)",1]
["B.18.6","Report Designer",1]
["B.18.7","Predicted Value generation",1]
["B.18.8","Language maker",1]
["B.18.9","User parameters",1]
["B.18.10","User predicted values",1]
["B.18.11","Network data base",1]
["B.18.12","Interface",1]
["B.19","Components:"]
["B.19.1","Operator guidance including quality control",1]
["B.19.2","Clearly arranged examination results",1]
["B.19.3","System concept",1]
["B.19.4","Highest level of patient safety and measurement accuracy ensured through build-in, automated feedback.",1]
["B.19.5","Should be suitable for use in all age groups",1]
["C.1","TOTAL WEIGHTED SCORE",32]

{"sheet":"Sheet1","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"hidden":true}